# Guideline | Canberra Health Services

## Cleaning of non-critical reusable medical devices (equipment, not instruments) and other patient care equipment

CHS26/084

Contents

Guideline statement	2

Scope	2

Section 1 - Risk assessment and risk rating	2

Section 2 - Types of non-critical RMDs, patient care equipment, and cleaning schedules	4

Section 3 – Detergents and Chemical Disinfectants	9

Evaluation	10

Related policies, procedures, guidelines and legislation	10

References	11

Definition of terms	11

Search terms	13

#### Guideline statement

##### Key objective

The cleaning of non-critical reusable medical devices (RMDs) (equipment, not instruments) and other patient care equipment is the responsibility of all Canberra Health Services (CHS)  staff. CHS includes inpatient and outpatient facilities at Canberra Hospital, Clare Holland House (CHH), North Canberra Hospital (NCH), University of Canberra Hospital (UCH), and community-based services. RMDs and patient care equipment used for non-critical tasks must be cleaned (reprocessed) in accordance with AS/NZS 5369:2023 and the manufacturer’s instructions for use.

This guideline outlines the requirements for the cleaning of any non-critical RMDs and other patient care equipment that is used during the delivery of patient care.

These requirements assist in preventing and controlling the transmission of infectious agents between patients, healthcare workers, and the environment.

Back to Contents

#### Scope

Cleaning of non-critical RMDs and other patient care equipment is undertaken by CHS staff and contracted cleaning services. This guideline applies to all CHS staff and contracted cleaning staff who use non-critical RMDs (equipment, not instruments) and other patient care equipment, or staff who are responsible for the cleaning of these items.   Staff are required to follow the manufacturer’s instructions for use to ensure that the correct cleaning requirements are met, and that the RMD or patient care equipment is not damaged by inappropriate use of cleaning products.

Alert: These guidelines do not provide specific cleaning instructions for all non-critical RMDs (equipment, not instruments) or other patient care equipment, rather it provides working examples.

Alert: For reusable medical devices that require cleaning and disinfection and/or sterilisation, please refer to the Reprocessing of Reusable Medical Devices Policy available on the CHS Policy and Guidance Documents Register.

Back to Contents

#### Section 1 - Risk assessment and risk rating

##### What are the risks

Any infectious agents (e.g. bacteria or viruses) introduced into the body can cause infection. In all healthcare settings, RMDs must be handled in a manner that prevents patients, healthcare workers, and the environment from coming into contact with potentially infectious material.

###### Principles for cleaning non-critical RMDs

Only Therapeutic Goods Administration (TGA)-approved RMDs should be used.

Before purchase, the manufacturer’s instructions for use must be reviewed by the local infection prevention and control team to ensure cleaning is compatible and achievable.

All RMDs used in the clinical environment must be reprocessed, including cleaning as per AS/NSZ5369:2023 according to their intended use and the manufacturer’s guidelines.

Single-use medical devices must not be reprocessed/cleaned or reused.

###### Principles for cleaning patient care equipment

Items such as computers and portable mobile devices (including rovers) and furniture and fixtures (e.g. over-bed tables and curtains) used in patient care, are classified as patient care equipment.

These types of items may not require TGA approval; however, the manufacturer’s instructions for use must be reviewed by the local infection prevention and control team prior to purchase.

Patient care equipment used in the clinical environment must be cleaned following use, and as per the manufacturer's instructions for use.

##### Assessing the degree of risk

This guideline provides advice on non-critical RMDs (equipment, not instruments) that are designed to be reused and require cleaning at the point of care, immediately after use. Risk is assessed using the Spaulding Classification Scheme and Table 1 outlines the different categories of risk.

Spaulding’s Classification

| Category | Description |
| --- | --- |
| Critical | These items confer a high risk for infection if they are contaminated with any microorganism and must be sterile at the time of use. This includes any objects that enter sterile tissue or the vascular system, because any microbial contamination could transmit disease. |
| Semi-critical | These items come into contact with mucous membranes or non-intact skin and should be single use or sterilised after each use. If sterilisation is not possible, high-level disinfection is the minimum level of reprocessing that is acceptable. |
| Non-critical | These items come into contact with intact skin but not mucous membranes. Cleaning is sufficient for most non-critical items, although intermediate or low-level disinfection may be appropriate in specific circumstances. |

Alert: Items which are reusable and come into contact with the patient or the patient zone, must be cleaned immediately after use and before use on the next patient (for example, observation machines or stethoscopes).

Back to Contents

#### Section 2 - Types of non-critical RMDs, patient care equipment, and cleaning schedules

This section provides an overview of commonly used non-critical RMDs. The details in this section are not exhaustive, and staff should apply the above risk rating and risk assessment before using any item when delivering patient care.  The following principles apply to all non-critical RMDs and patient care equipment, used in the delivery of patient care:

Any non-critical reusable equipment that comes into contact with a patient or their environment must be cleaned after use.

For any non-critical reusable equipment, where the staff member is uncertain if the item is clean, the staff member must clean the device/equipment before it is used in the delivery of patient care.

Adhesive tapes, labels or Blu-tack leave sticky residue and must not be applied to any reusable medical devices or patient care equipment.

If equipment is damaged, cleaning becomes ineffective. The equipment should be repaired or replaced.

The manufacturer’s instructions for use must be available to staff to ensure the correct cleaning procedure is being followed.

Only CHS supplied (TGA-approved) cleaning products are to be used.

Cleaning of the equipment is still required if a cover has been used e.g. temperature probe cover.

Where possible, after cleaning, equipment should be labelled as clean using a commercially available label which does not leave a residue.

After cleaning and where possible, a cover should be used to ensure the equipment remains clean.

Additional cleaning may be required during an outbreak. Procedures will be determined in consultation with the local infection prevention and control team.

Alert: The method of cleaning for individual items is not referenced in this guideline and staff should review and adhere to the manufacturer's instructions for use.

##### Common non-critical RMDs and patient care equipment

Table 2 provides a list of items which are commonly used during patient care and must be cleaned at the point of care by the person who used the device or equipment.

Common patient care equipment which requires cleaning at the point of care by the user

| Equipment Item | Frequency | Comment |
| --- | --- | --- |
| BGL Machine | After use | Clean protective case and the BGL machine |
| Breast Pumps | After Use |  |
| Chair (dental) | After use |  |
| Clipboard, pens and patient notes folder | Clean between patients |  |
| Commodes and shower chairs | After use |  |
| Computer on Wheels (COWs) and Rovers | Clean between patients | Keyboard covers or washable keyboards |
| ECG Machines | After use |  |
| Lead Gowns/Aprons | Between patients and after use |  |
| Mattress | Clean when visibly soiled or on discharge | Replace if waterproof cover is damaged |
| Over bed tray table | Twice daily |  |
| Patient Warmers | After use | Covers are single patient use |
| Pans and Urinals | After Use | Utilise pan flushers and sanitiser |
| Pillow (waterproof cover) | Clean when visibly soiled or on discharge | Replace if waterproof cover is damaged |
| Reusable injection trays | After and before use | Staff must adhere to the CHS aseptic technique procedure |
| Spine bath | After use |  |
| Telephone | After use |  |
| Torniquet (reusable) | After use |  |
| Trolley, dressing | Before and after use |  |
| Trolley, resuscitation | Daily and after use |  |
| Wash bowls | After use | Utilise washers |

Source: Australian Guidelines for the prevention and Control of Infection in Healthcare (2019).

##### Specialised equipment

Table 3 contains a list of items which are maintained by Healthcare Technology Management (HTM)/Biomedical Engineering. Although HTM and Biomedical Engineering are responsible for servicing and maintenance, staff using these items in the clinical area are responsible for cleaning them. All equipment must be cleaned prior to being sent for service or repair, and again on return from repairs.

HTM/Biomedical Engineering equipment and cleaning requirements by the staff at point of care

| Item | Frequency of cleaning | Comments |
| --- | --- | --- |
| Mechanical ventilator and non-invasive ventilation machines | Each shift and after use |  |
| Haemodialysis machines | After use |  |
| High flow nasal prongs | Daily and after use |  |
| Infusion pumps ( fluid, feeding etc) | Daily and after use |  |
| Vital signs equipment | After use and/or between patients | Staff should consider disposable BP cuffs for patients with a muti-resistant organism (e.g. MRSA) |
| Bladder scanner | Before and after patient use |  |
| Cardiac Monitors | Each shift and on discharge |  |
| Spirometer | After use and/or between patients | Mouthpieces are single patient use |
| Cardiotocography (CTG)Machines | After use and/or between patients |  |

##### Patient Physical Mobility Aids

Common physical mobility items and cleaning requirements

| Item | Frequency of cleaning | Comments |
| --- | --- | --- |
| Wheelchairs | After use | NOTE: equipment can be labelled and allocated to individual patients for their stay |
| Walking Frames | After use |  |
| Walking sticks | After use |  |

##### Manual Handling Equipment

Common manual handling equipment and cleaning requirements

| Item | Frequency of cleaning | Comments |
| --- | --- | --- |
| Patient slide sheets |  | Are single patient use and should be discarded if soiled with blood or bodily fluid |
| Hoist (frame) | After each use |  |
| Fabric hoist slings | Laundered between each patient via Capital Linen | The fabric sling may be dedicated to one patient for the duration of their stay |

##### Rehabilitation and Gym Equipment

Common rehabilitation and gym equipment

| Item | Frequency of cleaning | Comments |
| --- | --- | --- |
| Treadmill | Clean areas touched by patient after use Daily cleaning | Ask patients to perform hand hygiene prior to using equipment |
| Exercise Bike | Clean areas touched by patient after use Daily cleaning |  |
| Free Weights | Clean areas touched by patient after use Daily cleaning |  |
| Other Gym equipment | Clean areas touched by patient after use Daily cleaning |  |

##### Routine environmental cleaning: Patient room and shared/common spaces

At CHS routine environmental cleaning is undertaken with detergent and disinfectant. Detergents help to loosen the germs, whilst the addition of a disinfectant ensures that surfaces are cleaned and disinfected.

The CHS contracts manager and hospital assistant’s operations manager should be contacted for further details regarding the scope of practice for staff responsible for cleaning across CHS.

Table 7 lists items located within a patient’s room or in common areas across CHS sites, along with the required cleaning frequency for each.

Patient room and common areas

| Equipment Item | Frequency |
| --- | --- |
| Bath | Between use |
| Beds, Cots, Isolletes and Stryker trolleys | Clean frame or outside of isollete daily Clean whole bed on discharge or once per week for long stay patients |
| Call bell | Daily And On discharge |
| Ceiling | Spot clean as required |
| Chair (patient bedside) | Twice daily |
| Chair (common areas) | Daily |
| Computer (non-patient areas) | Daily or when visibly soiled |
| Curtains and blinds | Bed curtains - replace annually OR if visibly soiled Window blinds - spot clean as required. Clean annually |
| Doorknob/handle | Twice daily |
| Floor, non-slip | Damp mop twice daily and as required |
| Light switch | Daily |
| Linen skip (frames) | Daily |
| Mattress | When visibly soiled AND On discharge |
| Patient bedside locker / drawers | Daily AND On discharge |
| Medical gas equipment | Daily AND On discharge |
| Over bed tray table (over way table) | Twice daily |
| Sharps bin trolley | Weekly |
| Shower | Daily and as required AND For shared bathrooms - 3 times daily |
| Sink (hand wash basin) | Twice daily |
| Surfaces (general horizontal) in patient room e.g. ledges | Twice daily & spot clean after use |
| Toilet (patient) | Twice daily AND For shared bathrooms - 3 times daily AND On discharge |
| Televisions | Weekly AND On discharge |
| Walls (General) | Spot clean daily (includes dust & cobweb removal). Full clean 3 monthly |
| Walls (within operating/procedure rooms) | Clean after last procedure |

Back to Contents

#### Section 3 – Detergents and Chemical Disinfectants

##### 3.1 Neutral Detergents

The use of a neutral detergent solution or wipes on shared equipment, such as COWs, desktop computers, phones and workstations, is generally sufficient for cleaning, if the cleaning process undertaken follows the manufacturer’s instructions for use.

CHS utilise a range of neutral detergent wipes. Staff should always read the manufacturer’s instructions for use to ensure the wipe is compatible with equipment, prior to cleaning.

##### 3.2 Chemical Disinfectants

Only disinfectants specified on the Australian Register of Therapeutic Goods (ARTG) may be used by staff for disinfection, and the disinfectant must only be used for the approved purpose.

CHS  utilise a range of  2-in-1detergent disinfectant impregnated wipes. Staff should always read the manufacturer’s instructions for use to ensure the wipe is compatible with the equipment, prior to cleaning.

##### 3.3 Peracetic Acid (sporicidal wipes)

Peracetic acid-generating wipes are used for surface disinfection and for cleaning non-invasive medical devices. They are effective against Clostridioides difficile infection in two minutes. Staff should use this product instead of a standard  2-in-1 detergent disinfectant product when cleaning items used for patients with confirmed of suspected Clostridioides difficile infection.

Back to Contents

#### Evaluation

##### Outcome

Items and environments described in this guideline are cleaned as per recommendations

Reduction in healthcare associated infections

##### Measures

Review of environmental audit results to demonstrate cleaning is occurring as per procedure

Review of Multi-resistant organism, healthcare associated infection data rates for CHS.

Back to Contents

#### Related policies, procedures, guidelines and legislation

##### Policies

Waste Management

Reprocessing of Reusable Medical Devices

Introduction of New Health Technology

##### Procedures

Infection Prevention and Control in Healthcare Settings

Infectious Disease Outbreak Management

Infection Prevention and Control risk assessment and action plan for all construction and maintenance activities

##### Legislation

Health Records (Privacy and Access) Act 1997

Human Rights Act 2004

Public Health Act 1997

Work Health and Safety Act 2011

Australian Charter of Healthcare Rights

Waste Management and Resource Recovery Act 2016

Medicines, Poisons and Therapeutic Goods Regulation 2008

##### Standards

AS5369:2023 Reprocessing of Reusable Medical Devices in Health Service Organisations

Back to Contents

#### References

- Australian Commission on Safety and Quality in Health Care: Core Strategies for VRE prevention and control (2019)

- Australian Guidelines for the Prevention and Control of Infection in Healthcare, Canberra: National Health and Medical Research Council (2019)

- Clinical Excellence Commission: Infection Prevention and Control Practice Handbook (2020)

- Back to Contents

#### Definition of terms

| Term | Definition |
| --- | --- |
| Cleaning | The physical removal of soil and reduction in the number of microorganisms from a surface by a process such as washing in detergent. In order to decontaminate patient care equipment effectively, all organic debris (for example, blood, tissue and other body fluids) must be removed from the item during the cleaning process. |
| Detergent wipe | A detergent-containing wipe |
| Discharge Clean | Room/bed zone are cleaned as soon as possible when a patient is discharged or moves to another room. All patient furniture, bathroom, patient care equipment and accessories (e.g. nurse call bell) are cleaned by ISS and HAs |
| Disinfectant Wipe | The wipe is impregnated with a disinfectant |
| Disinfection | There are three levels of disinfection: - High-level disinfection: A disinfectant that kills all microbial pathogens, except large numbers of bacterial endospores when used as recommended by manufacturer. High-level disinfection is the minimum grade disinfectant used for disinfection of a semi-critical reusable medical devices (instruments and equipment). - Intermediate-level disinfection: A disinfectant that kills all microbial pathogens except bacterial endospores, when used as recommended by the manufacturer. This process is effective against most bacterial, viral and fungal spores (source: AS 5369:2023). - Low-level disinfection: A disinfectant that rapidly kills most vegetative bacteria as well as medium sized lipid containing viruses, when used as recommended by the manufacturer. It cannot be relied upon to destroy, within a practical period, bacterial endospores, mycobacteria, fungi or all small nonlipid viruses. Intermediate and low-level disinfection is the minimum grade disinfectant used for disinfection of a non-critical reusable medical or other devices where required. |
| Indirect contact (transmission) | Involves transfer of an infectious agent through a contaminated intermediate object or person e.g. hands of a HCW contaminated by contact with equipment/items from one patient area, with contact to another patient or person. |
| Non-critical equipment | These items come into contact with intact skin but not mucous membranes. Thorough cleaning is sufficient for most non-critical items after each individual use, although either intermediate or low-level disinfection may be appropriate in specific circumstances |
| Critical equipment | These items confer a high risk for infection if they are contaminated with any microorganism and must be sterile at the time of use. This includes any objects that enter sterile tissue or the vascular system, because any microbial contamination could transmit disease. |
| Semi-critical equipment | The items come into contact with mucous membranes or non-intact skin and should be single use or sterilised after each use. If this is not possible, high-level disinfection is the minimum level of reprocessing that is acceptable. |
| Standard Precautions | Involves the use of safe work practices and protective barriers; and are designed to reduce the risk of transmission of micro-organisms from both recognised and unrecognised sources of infections in health organisations. |
| Sterilisation | Sterilisation destroys all microorganisms on the surface of an instrument or device; to prevent disease transmission, details are available in Standard AS 5369:2023. |
| Shared patient equipment | Shared patient equipment is equipment used on more than one patient. |
| Therapeutic Goods | In most cases, therapeutic goods must be entered in the Australian Register of Therapeutic Goods (ARTG) before they can be lawfully supplied in Australia |
| Reusable Medical Device | Reusable medical devices (RMDs) are used for diagnostic and/or treatment purposes for multiple patients and must be approved by the TGA for reprocessing and reuse |
| Patient Zone | The patient zone is defined as any area where a patient can touch or have personal items, and it is considered contaminated. |

Back to Contents

#### Search terms

Infection Control, Cleaning, Shared Patient Equipment, Reusable Medical Device

- Back to Contents

For Policy Team to complete:

| Date amended | Section amended | Divisional approval | Final approval |
| --- | --- | --- | --- |
| 01 October 2021 | Complete review | Cathy Burns, Acting ED-NMPSS | CHS Policy Committee |
| 20/02/2024 | Document updated to include NCH | Judy Ryall, NCH ED Nursing & Midwifery | CHS Policy Team |
| 4 December 2025 | Review date extended for four months to allow NCH collaboration and consultation | Kellie Lang, ED of Nursing & Midwifery | Jacqui Clissold, Director of Policy and Consumer Handouts |
| 30/03/2026 | Complete Review | Kellie Lang, ED of Nursing and Midwifery | CHS Policy and Consumer Handouts |

This document supersedes the following:

| Document number | Document name |
| --- | --- |
| CHS20/290 | Infection Control – Cleaning (Shared) Patient Care Equipment Guideline |
| CHS21/560 | Infection Control – Cleaning (Shared) Patient Care Equipment Guideline |

Disclaimer

This document has been developed by Canberra Health Services specifically for its own use.  Use of this document and any reliance on the information contained therein by any third party is at his or her own risk and Canberra Health Services assumes no responsibility whatsoever.

- Back to Contents

| Acknowledgement of Country Canberra Health Services acknowledges the Ngunnawal people as traditional custodians of the ACT and recognises any other people or families with connection to the lands of the ACT and region. We acknowledge and respect their continuing culture and contribution to the life of this region. © Australian Capital Territory, Canberra 2024 | Accessibility call (02) 5124 0000 Interpreter call 131 450 canberrahealthservices.act.gov.au/accessibility |
| --- | --- |