# Policy | Canberra Health Services

Reprocessing of Reusable Medical Devices

CHS25/412

Contents

Policy Statement	3

Purpose	3

Alerts	4

Scope	4

Roles and Responsibilities	5

Section 1- General Principles for Reprocessing of Reusable Medical Devices	6

Section 2 – Cleaning of reusable medical devices	12

Section 3 - Disinfection	15

Section 4 – Packing and Inspection of RMDs	16

Section 5 – Sterilisation	18

Section 6 - Safe transportation, handling and storage of RMDs	20

Section 7 - Reprocessing equipment	21

Section 8 - Routine monitoring and control of reprocessing equipment and processes	25

Section 9 - Traceability of reprocessed semi-critical and critical RMDs	28

Section 10 - Audit Program	29

Section 11 – Staff Education and Training	31

Section 12 - Ultrasound probes	32

Section 12.1 – Non-critical	34

Section 12.2 – Semi-critical	36

Section 12.3 – Critical	43

Section 12.4 - Ultrasound Gel Selection, Probe Covers and Probe Storage	43

Section 12.5 – Acquisition of new Ultrasound machines, probes and consumables	45

Evaluation of the policy and processes	46

Related policies, procedures, guidelines and legislation	46

References	47

Definition of terms	49

Search terms	50

Attachments	50

Attachment 1 – Reprocessing of Ultrasound Probes at Point of Care Internal Audit Form	52

Attachment 2 – High Level Disinfection of Ultrasound Probes using Chronos Competency Assessment	56

Attachment 3 – Examples of Ultrasound Cleaning and Reprocessing	59

#### Policy Statement

This policy outlines the requirements for reprocessing (cleaning, disinfection and/or sterilisation) of reusable medical devices (RMDs) to ensure they are safe for use and minimise the risk of transmission of infectious agents. This includes all RMDs used in patient care areas within Canberra Health Services (CHS). The requirements of this policy must be met to provide a cleaned, disinfected and/or sterilised reusable medical device for patient use.

This policy should be read in conjunction with the relevant national and international standards, including:

Australian reprocessing standards AS 5369:2023 Reprocessing of reusable medical devices and other devices in health and non-health related facilities (AS 5369) and its normative references from clause 1.3 of the standard

manufacturers’ instructions for use (IFU)

Australian Guidelines for the Prevention and Control of Infection in Healthcare

Infection Control in Endoscopy Guidelines 2025 (GENCA/GESA)

relevant guidelines for reprocessing of ultrasound transducers:

  - ASUM - Australasian Society for Ultrasound in Medicine

  - AJUM - Australasian Journal of Ultrasound in Medicine

National Safety and Quality Health Service (NSQHS) Standards

Australasian Health Facility Guidelines - 0190 – Sterilizing Services and Endoscope Reprocessing Units

#### Purpose

The purpose of this document is to outline:

governance systems in place at CHS for reprocessing RMDs

CHS compliance with the requirements of:

AS 5369

Australian Guidelines for the Prevention and Control of Infection in Healthcare

Infection Control in Endoscopy Guidelines 2025 (GENCA/GESA)

manufacturer’s IFU

system and processes for safe and adequate reprocessing of RMDs in sterilising services

management and reprocessing of ultrasound probes at point of care by the end users to minimise the risk of transmission of microorganisms before, during, and after ultrasound use.

This policy excludes RMDs/other devices used in any of the following activities:

autopsy, post-mortem examinations, anatomy dissections of deceased persons, mortuary care and embalming

veterinary procedures

Preclinical training using animal or removed human tissue (e.g. extracted teeth).

Back to Contents

#### Alerts

AS 5369 does not provide reprocessing advice for:

medical devices that are intended for single use only

reprocessing of RMDs potentially contaminated with Creutzfeldt-Jakob Disease (CJD) CJD recommendations are described in the Creutzfeldt-Jakob Disease (CJD) Transmission Guideline available on the CHS Policy and Guidance Documents Register

the sterilisation of bandages, dressings, and other consumable products in facilities

items not intended to be used as medical devices

Staff who are required to perform automated methods of High Level Disinfection (HLD) of ultrasound probes at point of care must complete the relevant competency training on how to use the Automated High Level Disinfectors, refer to Section 2.2.

Back to Contents

#### Scope

This policy applies to the staff undertaking reprocessing of RMD at CHS including:

staff working in the central sterilising unit (Sterilising services)

North Canberra Hospital (NCH) Sterilising staff performing pre-cleaning prior to transporting the instruments to sterilising services unit located in building 5, Canberra Hospital (CH).

healthcare workers who perform pretreatment (bedside cleaning i.e remove gross soil using sterile water and wipe or sponge to prevent drying of contaminants) or pre-cleaning or cleaning or high level disinfection or oversee storage of reusable medical devices at point of care within CHS including:

Perioperative Unit – perform pre-treatment of surgical RMDs, high-level disinfection of ultrasound probes at the point of care and oversee storage of sterile RMDs

Gastroenterology and Hepatology Unit – perform cleaning and high- level disinfection of flexible endoscopes and non-channelled scopes

Ward areas – perform pre-cleaning of instruments at point of care prior to returning to central sterilising unit, high level disinfection of ultrasound probes at point of care and manage storage of reprocessed RMDs

Healthcare workers who use ultrasound probes/machines, including bladder scanners, dopplers and handheld point of care ultrasounds

This document applies to the following staff working within their scope of practice and who are responsible for cleaning ultrasound probes at point of care:

Medical Officers

Nurses and Midwives

Allied health professionals

Sonographers

Equipment Officers

Hospital Assistants

Students working under direct supervision

Back to Contents

#### Roles and Responsibilities

CHS staff involved in reprocessing must adhere to this policy where applicable and the individual area managers or supervisors responsible for implementing the relevant applicable sections and monitor compliance with the NSQHS standards.

##### Sterilising Services

Sterilising Services provides comprehensive reprocessing of reusable medical devices including cleaning, disinfection, packaging, inspection & testing, sterilisation, storage, and transport services in the following sites:

Canberra Health Services Sterilising Services in Building 5

NCH Sterilising Services (pre-cleaning facility only)

Central Reprocessing Unit (CRU) in building 1 within Canberra Hospital

Sterilising Services is responsible for providing services to:

Perioperative and Interventional services

External stakeholders (private clinics)

Clinical areas including (but not exclusive to) Medical Imaging, Cardiac Catheter Laboratory, Oncology, Intensive Care Unit (ICU), Sleep and Respiratory Laboratory, Women and Children Maternity and Delivery suite, Central Outpatients, Paediatrics, Gastroenterology and Hepatology unit.

Sterilising services management and staff are responsible for ensuring RMDs are adequately cleaned and safe for use.

##### Sterilising Services management and staff ensure that the following are effectively managed and performed:

ensure appropriate governance systems are in place to facilitate reprocessing of reusable medical devices in accordance with the relevant national and international standards, and guidelines

ensure sufficient resources are in place to enable effective systems for safe reprocessing of reusable medical devices in accordance with the relevant national and international standards, and guidelines.

ensure that the day-to-day responsibility for establishing and monitoring the implementation of this policy is delegated to the relevant managers and staff.

manage human resources and co-coordinate day-to-day management of all areas within Sterilising Services in accordance with policies and procedures.

provide leadership in accomplishing the requirements of the quality management system, accreditation, system improvements and audits in accordance with policies and procedures and national standards.

manage the team to ensure the processes and education are in accordance with the national standards, quality management system (QMS), and applicable legislations (e.g. procurement and financial management), policies and procedures.

ensure that the new RMDs are in line with manufacturers’ specifications and the required documentation is provided prior to reprocessing and meet AS 5369 or equivalent standards and prepare reports for applicable forums.

identify and report on any deficiencies in operation of production area and surgical item inventory any deficiencies in quality standards to the SS management and recommend appropriate remedial action.

provide and oversee training for new and existing staff across sterilising services including timely completion of in-services and mandatory training and competency assessments.

ensure the training documentation is up to date to align with the current requirements of the standards for safe and effective reprocessing.

##### Staff reprocessing ultrasound probes at Point of Care

are trained and competent in pre-cleaning, cleaning, and high-level disinfection of the ultrasound probes. Refer to Section 2.2.

reprocess ultrasound probes according to this policy and relevant standards.

- Back to Contents

#### Section 1- General Principles for Reprocessing of Reusable Medical Devices

RMDs are used to diagnose and treat multiple patients and are intended by the manufacturer for reprocessing and reuse. Examples of RMDs include surgical forceps, endoscopes, laryngoscopes and ultrasound probes (transducers). Reprocessing is a multistep process, and the requirements of the AS 5369 shall be adhered to:

all RMDs are reprocessed according to their intended use and in accordance with the manufacturer’s instructions and requirements stated in the ISO 17664-1 and ISO 17664-2 standard

only Therapeutic Goods Administration (TGA) registered RMDs, cleaning agents, disinfectants or sterilants are used (TGO 104)

- policies, work instructions, audit checklists, routine monitoring forms and procedures for the reprocessing activities are documented and dated in the QMS according to Section 2.2.2 in AS 5369

- reprocessing activities of RMDs are performed in designated reprocessing areas

- standard precautions from the Australian Guidelines for the Prevention and Control of Infection in Healthcare are applied when cleaning and handling used RMDs, additional transmission based precautions apply as required

- appropriate personal protective equipment (PPE) shall be worn in the reprocessing area (scrubs, hair net, enclosed shoes, gloves when handling dirty RMDs, waterproof thumbs up gown, disposable face shield)

- RMD manufacturer’s instructions, TGA certificate and completed provisional form from the supplier are reviewed and approved by sterilising services prior to purchasing to ensure they comply with the existing reprocessing methods and equipment. RMDs that don’t comply can’t be purchased

- damaged RMDs are reported by the end users and by the sterilising team. The instrument manager sends damaged RMDs out for repair or a replacement RMD is sought as required

- following comprehensive in house training, Sterilising Services staff complete competency based assessments prior to performing undertaking reprocessing activities with limited supervision. The Training and Development Manager or Team leaders assess staff as competent and the time period between reassessment of competency is 12 months from the date completed

- When configuring an RMD into sets they should weigh up to 5 kilograms and not exceed 7 kilograms (refer to NSW Workcover Authority (2011), Design and handling of surgical instrument transport cases: A guide on Health and Safety Standards)

Note: In exceptional cases only, if an RMD manufacturer’s instructions are slightly different to the validated processes in sterilising services then sterilising services in consultation with the manufacturer/supplier will undertake risk assessment and can choose to deviate if the risk involved is minimal. The  RMD is validated in sterilising services using the alternative process where required (Risk analysis is stipulated in Appendix B of AS 5369)

##### Spaulding Classification Scheme

As defined in AS 5369, the Spaulding Classification Scheme is used to categorize an RMD according to its intended use and the subsequent level of reprocessing required to render the RMD safe for reuse.

Critical - RMDs require cleaning followed by sterilisation (comes in contact with sterile tissue or cavity)

Semi-critical - RMDs require cleaning followed by high-level disinfection (comes in contact with non-intact skin or mucous membrane)

Non-critical - RMDs require cleaning followed by low or intermediate level disinfection (comes in contact with patient’s intact skin)

Spaulding Classification Scheme

| Classification/ Level of risk | Application | Process | Storage | Examples |
| --- | --- | --- | --- | --- |
| Critical | An RMD that comes into contact with sterile tissue, cavity or bloodstream shall be sterile at the time of use | Clean as soon as possible after use. Sterilisation by moist heat e.g. Steam under pressure. If RMD is heat or moisture sensitive sterilise using alternative process e.g. Ethylene Oxide, low temperature chemical sterilising systems (e.g. hydrogen peroxide) or with a liquid chemical sterilant. | Sterility should be maintained: Packaged RMDs should be dry The integrity of the wrap should be maintained Wraps should act as an effective bio barrier during storage Store to protect from environmental contamination in designated storage area Unpacked RMDs should be used immediately | RMDs used for invasive procedures Needle Guides Biopsy Needles (reusable) BIP (biopsy) guns Sentinel Node Probes Rigid Endoscopes and accessories used in invasive procedures including Arthroscopes Hysteroscopes Laparoscopes Rigid Bronchoscopes Flexible endoscopes such as Cystoscopes |
| Semi-critical | Contact with intact nonsterile mucosa or non-intact skin | Clean as soon as possible after using. Steam sterilisation is preferred where possible, or high-level thermal disinfection i.e. Heat. If the equipment will not tolerate steam sterilisation or high-level thermal disinfection, use high-level chemical disinfection or low temperature automated chemical sterilising systems. | Store to protect from environmental contamination in a designated storage area to protect RMDs. Store to protect from environmental contamination. | Respiratory support devices e.g. breathing circuits, air vivas Vaginal speculae Invasive ultrasound probes Transoesophageal echocardiogram ultrasound (TOES) probes Transrectal ultrasound (TRUS) probes Transvaginal (TV) ultrasound probes Flexible endoscopes Colonoscopes, Gastroscopes & Sigmoidoscopes |
| Non-critical | Contact with intact skin | Clean as soon as possible after using. Clean with detergent and water. If required, disinfect with a compatible low/intermediate level instrument grade disinfectant after cleaning. | Store in a clean, dry place. | Non-invasive devices e.g. abdominal ultrasound or perineal probes used on intact skin and head clamp system Bedpans and urinals Blood pressure cuffs Linens, beds Mercury thermometers Stethoscopes Sphygmomanometers & tourniquets |

##### Product families

Product families are groups of RMDs/other devices which share similar attributes (eg. mass, design, and construction materials, and which present a similar challenge to cleaning, disinfection and sterilisation processes). Sterilising Services assigns RMDs into product families based on the intended use of the device, the materials of construction, design of, physical characteristics and packaging of the device. Product family ensures that the RMD is being reprocessed at the right temperature and time.

When assigning an RMD to a product family and to a method of reprocessing, the following is considered and documented:

- the range of steam penetration resistance

- description of the RMD and its intended use with reference to the IFU and Spaulding classification scheme

- description of materials that are used to make the RMD. Note: both metal, non-metal and metal non-metal combinations are commonly used to construct.

- RMDs and materials can affect penetration by the reprocessing agent(s)

- the design of the RMD, including characteristics that could affect the selection of a cleaning, disinfection, or sterilisation process, e.g. ease of disassembly and assembly, tolerance to moisture, heat and chemicals, presence of lumens, moving parts, fibre-optics, and electronics

- the physical characteristics of the RMD, including its mass, surface area and thermal conductivity

- packaging of the RMD, including the Sterile Barrier System (SBS) for sterilised devices. The packaging system should be compatible with the sterilisation process. The final load mix with the packaging systems should also be considered with product family categorisation as the final load configuration (with inclusion of packaging systems).

Once the above items are identified, validation tests on the most challenging device or group of devices for each process to identify the master device for that process is undertaken. The master devices are identified and used to conduct performance qualification taking into account the SBS and the device when processed together in a set. Further information is available in ISO standards 17665 and ISO 17664-1.

Back to Contents

#### Section 2 – Cleaning of reusable medical devices

##### Cleaning Agents and Disinfectants

Certified and compliant cleaning agents and instrument chemical disinfection systems (high-level disinfection systems) are in place. Cleaning agents, disinfectants and sterilising agents are compatible with the RMDs to be processed, and the associated equipment used to deliver that process. Chemical disinfectants can be inactivated by organic material; therefore, items for disinfection are thoroughly cleaned prior to disinfection. A cleaning agent is used to remove residual soil/organic matter from a used RMD.

Documented specifications is obtained for each cleaning agent that provides the following requisite information as outlined in AS 5369, clause 3.1.3.

safety data sheet and regulatory status

active ingredient(s) and physical/chemical properties, including stability (shelf life)

microbial efficacy and toxicity/residues

material effects of the agent on RMDs

compatibilities and known device material non-compatibilities.

container/packaging/labelling (including shelf life and storage requirements)

directions for use and, where intended for the product, reuse

Cleaning agents are:

intended for use on medical devices and used only if they are on the Australian Register of Therapeutic Goods (ARTG)

compatible with the RMDs being processed and the selected method of cleaning.

diluted and used in accordance with their IFU.

compatible with the available water quality, and preferably biodegradable

non-toxic, non-abrasive, low foaming, free rinsing, and preferably in liquid form

Workflow and processes for bedside pre-cleaning: At point of use, the end users should use sterile water and a lint free cloth to remove gross soil from the used RMDs. This is to minimise the risk of drying of contaminants prior to the full reprocessing process.

As recommended in the AS 5369 standard, ACORN (Standards for perioperative nursing in Australia), Australasian Health Facility Guidelines HPU520 Operating Unit “Reusable medical devices should be wiped clean with a sponge after immediate use or during the procedure by nursing staff. In some orthopaedic cases, RMD are soaked during the case as the debris can be difficult to remove once it is left for a period of time.”

Sterilising services has a comprehensive system for cleaning of used RMDs.  The process is reviewed regularly to ensure that the first stage of reprocessing mitigates risks associated with incorrectly reprocessed RMDs.

A delay between the initial pre-treatment and the validated cleaning process for that device can result in an increase in the bioburden (through microbial proliferation). It can also cause adherent material to dry on the device making removal of this material more difficult.

Before returning RMDs to the reprocessing area/central sterilising unit, all visible organic material/gross soil is removed by the end users at point of use i.e., removal of visible blood, bone, tissue and debris intra-operatively or immediately post-operatively from the RMD. The RMDs are initially pretreated at the bedside /point of care prior to return to the reprocessing area. In the perioperative unit, it recommended that RMDs should be wiped clean with a sponge and sterile water after use by nursing staff. In some orthopaedic cases, RMDs are soaked during the case as the debris can dry and be difficult to remove once it is left for a period of time.

All RMDs are placed in sealed tubs or enclosed case carts for transportation, refer to section on transport.

Cleaning always precedes disinfection and sterilisation and incorporates the following practices:

be performed as soon as possible after use. However, some RMD manufacturer’s recommend six hours as the maximum time lapse between bedside cleaning and full cleaning in sterilising services, ideally as soon as the procedure ends, preferably within an hour post-surgery

For flexible endoscopes, the time lapse is 15 minutes preferably or within an hour. Where delays in the validated cleaning process occur after surgery, a risk assessment is undertaken and manufacturer’s instructions are referred to for delayed processing. Gross soil is wiped down from the flexible endoscope with a damp cloth. Aspirate sterile water or detergent through the channel and is flushed

RMDs are inspected when returned to the SS (damaged/missing are reported to the perioperative unit team leader or patient flow coordinator or resource manager) and the decision to use the tray without the instrument is advised by the perioperative unit

disassembly or assembly as per manufacturer’s IFU before cleaning, disinfection, or sterilisation

a manual or automated process is followed as per the IFU (manual cleaning may be specified by the manufacturer’s IFU prior to automated cleaning and/or disinfection or sterilisation processes)

cannulated RMDs are brushed, flushed, and rinsed prior to washing them in the washer disinfectors by connecting them to the additional accessories called the minimally invasive instrument set (MIS) baskets or trays

ultrasonic cleaning is performed only if it is mandatory as per the manufacturer’s instructions. Water and cleaning agent solutions are changed daily, and whenever soil is visible in the tank. Routine tests are performed to check the efficacy of the equipment

Heat labile RMDs are manually brushed and cleaned at the sink. Final rinsing after manual cleaning is performed with the reverse osmosis (RO) water. Brushes and other accessories used for manual cleaning are TGA certified. If reusable brushes are used, they are to be cleaned and thermally disinfected or sterilised at least daily.

Cleaning procedures include:

If a Washer Disinfector (WD) is available:

manual cleaning only

manual cleaning & ultrasonic cleaning & WD

cleaning in a WD directly

If WD is not available:

- manual cleaning only

manual cleaning and ultrasonic cleaning

Manual cleaning of RMDs is used only:

- if an RMD’s validated cleaning instructions require manual cleaning; or

- as a pre-treatment prior to reprocessing of an RMD in a WD

The ultrasonic cleaner cleaning unit is an additional accessory for cleaning RMDs. Powerful ultrasonic cavitation breaks apart gross soils and washes hard to reach areas that brushes, and sprayed water alone cannot. The standard cannulated instrument irrigator ports help in flushing of cannulated instruments during wash and rinse cycles. loosens the debris. The RMD undergoes a full validated cleaning process following the cavitation action.

An automated cleaning process in a WD is used because an automated process is easier to replicate than a manual cleaning process.

Documented cleaning procedures/ work instructions for cleaning of RMDs including flexible endoscopes (leak testing, manual cleaning and flushing, high level disinfection and/or sterilisation) are accessible on sterilising services QMS.

Flexible endoscopes are reprocessed in accordance with:

AS 5369

Infection Prevention and Control in Endoscopy 2025 Guidelines authored by the Gastroenterological Society of Australia (GESA)/Gastroenterological Nurses College of Australian (GENCA)

manufacturer’s IFU

Alert: RMDs used on procedures involving high risk tissues (brain; dura mater; pituitary gland; spinal cord; posterior segment of the eye; cranial and dorsal root ganglia and olfactory epithelium) are disinfected using the prion deactivator chemical and sterilised at 134 degrees Celsius for 18 minutes steam exposure. Accessories such as cleaning brushes used to clean RMDs used on high risk tissue are discarded after use. The prion chemical meets the requirements specified in the Creutzfeldt–Jakob disease (CJD) guidelines, including that the chemical is listed on the ARTG together with a specific statement confirming that the product is indicated for the removal or inactivation of prions. Use of this prion deactivating chemical, in conjunction with an 18-minute sterilising cycle, achieves removal/deactivation of prions, rendering the RMDs used on high risk tissues safe for subsequent patient use, however if the RMD has been used on the high risk tissue on a suspected CJD patient, the requirement to destroy, or quarantine RMDs would still have to be followed in consultation with the Perioperative Unit, Infection Prevention Control Unit (IPCU) and the CJD registry. RMDs that are used on low infectivity tissues are washed in the normal cycle using a neutral to mildly alkaline chemical.

Back to Contents

#### Section 3 - Disinfection

A high-level instrument grade disinfectant is the minimum grade disinfectant used for disinfection of a semi-critical RMDs. Where disinfection of a non-critical RMD is required, an intermediate or low-level instrument grade disinfectant shall be the minimum grade disinfectant used.

Disinfectants differ significantly in their spectrum of antimicrobial activity and their speed of action as follows:

low-level instrument grade disinfectants kill vegetative bacteria, some fungi, and viruses.

intermediate-level instrument grade disinfectants kill vegetative bacteria, mycobacteria, viruses, and most fungi but do not kill bacterial endospores.

high-level instrument grade disinfectants kill all microorganisms except for high numbers of bacterial endospores. Some disinfectants used as high-level instrument grade disinfectants are chemical sterilising agents that kill high numbers of bacterial endospores with prolonged exposure under controlled and defined conditions

Semi-critical RMDs that cannot withstand high temperature are disinfected using a chemical disinfection process. For chemical disinfection, the RMD is rinsed in accordance with specific manufacturer’s instructions to remove any residual detergent. Flexible endoscopes and other heat-labile semi-critical items are disinfected in the automated flexible endoscope preprocessors (AFERs) that follow the normative standard ISO 15883 parts 1 and 4 in conjunction with manufacturer’s IFU, specific requirements of Infection Prevention and Control Guidelines 2025 and AS 5369 requirements.

A chemical disinfectant used to reprocess an RMD is labelled as an instrument grade disinfectant and other classes of chemical disinfectants e.g., hospital grade disinfectant are not used to reprocess an RMD as these are intended for hard surfaces not medical devices. Skin disinfectants (e.g., chlorhexidine) are chemicals formulated for use on skin or tissue and must not be used to clean equipment.

Heat stable RMDs are disinfected using the thermal disinfection process. Prior to thermal disinfection, the RMD is thoroughly cleaned. Cleaning and thermal disinfection is performed in a washer disinfector (WD) compliant with the ISO 15883-3 normative standard. The minimum temperature and holding times to achieve thermal disinfection is 91 °C for 1 to 5 minutes.

For specific recommendations on disinfection type and method, the applicable manufacturer’s IFU and the Spaulding classification is followed.

Ultrasound probes and other semi critical devices are disinfected in the UVC (Ultraviolet light-C) HLD (Chronos) at point of use.

Refer to Section 2.1, 2.2 and 2.4 for more information on ultrasound probe disinfection.

Disinfected RMDs are labelled with sticker/ TDOC label with the below details:

date and time of processing

equipment information

process/cycle information

chemical and/or disinfectant information as applicable

set type and description of RMD

staff responsible for reprocessing

Back to Contents

#### Section 4 – Packing and Inspection of RMDs

##### Inspection, assembly and testing of RMDs prior to sterilisation.

Examination of an RMD prior to its assembly and packaging (including containment devices, such as instrument trays or instrument pins/stringers) includes the following checks:

surfaces, lumens, and grooves of the device are visibly free from soil, rust, or lint (where applicable, a style or a correct size clean, dry brush should be able to pass through the lumen)

all cutting edges of the device are sharp (where applicable, sharpness should be tested in accordance with the IFU)

all surfaces and non-cutting edges of the device are smooth, well finished, free of burrs and that there is no visible evidence of pitting or flaking

hinges, box joints and crevices are checked for ease of movement and are not damaged or cracked

jaws, teeth and tips of the device are not hooked or snagged and that they approximate accurately

serrations are uniform and not worn

ratchets close easily and hold firmly

all screws on a jointed device are present and are presented for sterilization as per IFU.

a device with valves moves freely and is left in the “on” or open position

a multi-component device is disassembled for cleaning in accordance with the IFU.

post cleaning the device is checked for completeness and correct function

telescopes and light cables are checked for dents, damage, and correct function in accordance with the IFU

an insulated device is examined and checked for integrity to confirm that the insulating layer is intact. The insulation test result is documented on the Sterilising Services electronic tracking system as evidence in the event of a recall or an incident

flexible lumen and non-lumen endoscopes are checked for completeness and function, e.g. attachment of accessories such as a sterilization cap according to the IFU

Power tools and battery operated RMDs are inspected for damage prior to packing

Only trained and competent staff whose competency assessment on reprocessing processes is up to date perform functional testing of an RMD as recommended on the IFU of the RMD and/or testing device. An RMD that is found on examination to be faulty or in need of repair is immediately removed from use and reported to the sterilising team leaders and the perioperative resource manager via an email or via the repair of surgical instruments reporting form. Lubrication of an RMD is performed in accordance with the device’s IFU. The lubricant is water miscible and compatible with the sterilising agent.

##### Packaging materials

RMDs are packed in the Sterile Barrier System (SBS)/ Preformed Sterile Barrier System (PSBS) wrapping material or peeled down pouches for terminal sterilisation of critical devices. This is the minimum package that minimises the risk of ingress of microorganisms and allows aseptic presentation.

The types of SBS used are:

sterilisation wraps

sealable pouches and reels

dust covers are used an additional protective layer where required

The type of SBS used for a specific RMD is determined by the type of sterilising process required for that RMD, noting that:

packaging and wrapping materials permit the removal of air from the pack, penetration of the sterilising agent and removal of the sterilising agent

trays are configured according to product families in relation to design, complexity, and reprocessing requirements

single-use wraps are discarded once used

The SBS / PSBS must be manufactured in compliance with ISO11607-1 and the applicable parts of EN 868 (series), Packaging materials and systems for medical devices which are to be sterilised.

Back to Contents

#### Section 5 – Sterilisation

Sterilisation destroys microorganisms on internal and external surfaces of RMDs/other devices, rendering them free from viable microorganisms. Sterilisation involves use of physical or chemical processes to destroy all microbiological life, including bacterial spores present on an RMD. Moist heat sterilisation is the preferred method for sterilisation of heat stable RMDs including its packaging. Where an item cannot withstand a moist heat sterilisation process, low temperature vaporised hydrogen peroxide sterilisation (VH2O2) is used.

Sterilising services provides terminal sterilisation to all critical RMDs to achieve the required sterility assurance level. The ISO standards for sterilisation processes are EN 285 (steam) and ISO/TS 22421 and 22441 (low temperature hydrogen peroxide).

Immediate use sterilisation:

| Type of process | Release criteria |
| --- | --- |
| Cleaning manually | RMD/other device is visually clean and dry. |
| Cleaning automated | RMD/other device is visually clean and dry. Cycle records conform to process specification. |
| Disinfection thermal | RMD/other device is visually clean and dry. Cycle records conform to process specification. |
| Disinfection chemical | Disinfection agent concentration, exposure time and temperature conform to process specification. RMD/other device has been rinsed in accordance with IFU. |
| Other high-level disinfection systems (not thermal or chemical) | RMDs/other devices that are compatible with the disinfection system are clean and any specific system controls and indicators confirm effectiveness of the process and expected performance of the equipment. |
| Packaging system — PSBS — Sealable pouches and reels | Packaging system is suitable for sterilisation process and of correct size for RMD/other device to be sterilized, is correctly labelled, the PSBS and seal are intact, CI is present with specified colour change (refer to ISO 11607 (series)). |
| Packaging system — SBS — Sterilisation wrap, i.e. woven and non-woven | Packaging system utilised is suitable for sterilisation process and of correct size for RMD/other device to be sterilised, is correctly labelled, packaging material and seal are intact, CI tape is present with specified colour change (refer to ISO 11607 (series)). |
| Sterilisation moist heat with drying cycle | Cycle records confirm achievement of process parameters established during PQ and external chemical indicators show specified colour change. Packaged RMDs are intact, there is no visible moisture. If used, results for BIs, PCDs, electronic data logger are correct (refer to ISO 17665 (series)). |
| Sterilisation peracetic acid | Cycle records confirm achievement of process parameters established during PQ and BIs and CIs show correct results (refer to ISO 14937). |
| Sterilisation vaporized hydrogen peroxide system | Cycle records confirm achievement of process parameters established during PQ and BIs and external CIs show correct results. Packaged items are intact, there is no visible moisture. If used, results for PCDs are correct (refer to ISO 22441). |
| Sterilisation — Other methods | Cycle records confirm achievement of process parameters established during PQ and external CIs show specified colour change. No growth detected from exposed BIs/PCDs. Packaged items are intact, there is no visible moisture. Records confirm external aeration (if required) performed in accordance with process specification. Reference may be required to IFU (refer to ISO 14937). |

Immediate use sterilisation is not routinely performed unless there has been an accident (instrument dropped on the floor) and is required immediately for the surgery. An agreement is in place between the perioperative services and sterilising services delegate to determine and approve immediate use sterilisation for an instrument in the event of an accident. The decision to hand wash instead of the recommended automated wash will be decided by the management at the time of the incident and the instrument will be sterilised through the shortest sterilisation cycle. This deviation will be documented both in sterilising tracking system and the patient records.

##### Release of RMDs after reprocessing

RMDs are not released after the final step of reprocessing until all acceptance criteria for release of the RMD have been met. If the below criteria specified in Table 9.1 of AS 5369 are not met, then the device is designated as nonconforming and will be investigated by sterilising services.

Cleaning, disinfection and sterilisation process cycle records are checked by the staff approving and releasing the load

Where used, the results of the cleaning indicators, biological indicators, chemical indicators and process challenge devices are checked

The traceability systems (T-DOC) are able to identify the type of RMD, serial number, date of reprocessing, identification of the competent person, identification of the reprocessing equipment, cycle numbers, date of reprocessing and other records, where required.

Back to Contents

#### Section 6 - Safe transportation, handling and storage of RMDs

RMD’s are handled, transported, and stored in accordance with clause A9.5 of AS 5369.

Incorrect handling of a reprocessed RMD during storage or transport of the RMD can compromise sterility and cause damage.

RMD’s sterilised by steam are allowed to cool prior to handling.

##### Safe transportation to and from the reprocessing - the collection of used RMDs from point of use

Training is provided on the collection and delivery of a used RMD/other device. Separate, designated containers and case carts are used for the collection and delivery of used RMDs to the Sterilising Service Unit. Selection of these containers consider the characteristics such as size of the devices to be transported. At a minimum, the case carts or enclosed tubs are:

- rigid, robust, and leak-proof

- of an adequate size to contain the device safely, to minimize hazards to staff handling the device and to prevent damage to the device during transport

- able to be securely sealed or locked to prevent tampering (where applicable)

- able to be clearly labelled to identify the hospital

- able to be readily cleaned, or cleaned and disinfected and dried

- able to be safely disposed of when no longer serviceable

Trolleys/case carts/tubs used for the transportation of used RMDs are cleaned with an appropriate detergent wipe between use.

Tubs or case carts used for transport of soiled / used RMDs are designated only for used RMDs and are cleaned and disinfected in the wash station or the cart washer between uses. TGA approved detergent wipes are used where washing of the transport containers is not possible.

##### Transportation of ultrasound probes

Disinfected ultrasounds probes are placed in clean, secured transport tubs for transportation. Green tubs are designated for disinfected probes (clean) and yellow for dirty probes. The tubs and trolleys are cleaned with TGA approved detergent wipes between use.

##### Handling of specialised RMDs including instruments on loan and trial

Loan equipment is formally checked in, cleaned, disinfected and sterilised by designated Loans Officers as per the current standards and guidelines. Loans reprocessing procedure is accessible in the sterilising services QMS (STR-GEN-PR-121).

##### Storage of RMD

The recommended acceptable range for storage of sterile RMDs is a temperature of 18 to 25 degrees Celsius, and relative humidity of 35% to 70%.

Flexible endoscopes that undergo high level disinfection are stored in validated controlled environment endoscope storage systems (plasma typhoon storage system and plasma bags) that meet the requirements of EN 16442.

Back to Contents

#### Section 7 - Reprocessing equipment

##### Washer Disinfectors (WD)

WDs are designed to clean and disinfect RMDs. WDs are preferred over the manual process if the IFUs permit. WDs clean the RMDs and remove blood, body fluids, debris, and other contamination on used RMDs. WDs disinfect RMDs based on the minimum requirements and comply with ISO 15883 applicable parts.

WDs are annually calibrated and validated with programmed cycles to ensure RMDs are adequately being reprocessed.

Sterilising technicians are to ensure that:

- the machine has been validated

- daily routine monitoring checks have been performed and passed

- physical integrity of internal chambers is intact and clean.

- racks used in WDs are checked for integrity and cleanliness.

- delicate or small items are placed in baskets with hold down catches

- devices with lumens are connected to the fluid pathway of the WD using appropriate adaptors / connectors

- cleaning efficacy tests are performed as recommended by the standards and the device manufacturer

##### Ultrasonic cleaners

Where an IFU specifies, the ultrasonic cleaners are used to remove gross soil where pre-treatment of a device requires use of an ultrasonic cleaner prior to manual and/or WD processing. The lid is closed when operating the ultrasonic cleaners to prevent the emission of aerosols during use. Cannulated RMDs are connected to the adapters for efficient cleaning. The performance of the ultrasonic cleaner shall be tested daily in accordance with AS 5369 and AS 2773:2019 (Clause 8.2.4).

##### Heat sealers

Heat sealers are used to seal film to paper (e.g., laminates, flexible packaging systems) and plastics by pressing the lacquered surfaces between heated plates. The temperature, pressure and contact times are constantly (daily) monitored, as creases, thickness and type of material used may result in faulty seals. Heat sealers used for sealing PSBS, or dust covers shall be operated in accordance with the IFU of the machine. Normative references ISO 11607-2, ISO TS 16775 have detailed information. The heat seals provided to the packs maintain integrity of terminally sterilised critical devices until point of use.

Key points:

seals are always checked on opening to ensure that the seal has been maintained.

on a daily basis, at least one sample of heat-sealed PSBS is checked for seal integrity before and after exposure to the sterilisation process.

adhesive tapes such as ‘sterilisation indicator sealing tape’ are used to fasten wrappings and incorporate a chemical indicator that changes colour during the sterilisation process.

T-DOC labels are attached prior to sterilisation

heat sealers undergo a complete mechanical service, including temperature calibration, at regular intervals not exceeding 12 months.

##### Sterilisers

Critical RMDs are terminally sterilised by steam or vaporised hydrogen peroxide sterilisers. The availability of safely reprocessed medical devices is provided by compliant sterilisers that conforms with the requirements of AS 5369.

The sterilisers are checked each day to ensure they are functioning as intended and the method of sterilisation selected for the RMD aligns with the manufacturer’s IFU. When loading the steriliser, the types of RMDs included in the sterilisation loads will be determined in accordance with the steriliser manufacturer’s IFU and the product family classification. Any restrictions or limitations relating to the size, mass, configuration, or loading orientation of devices being processed is considered during annual validation process and documented accordingly. The sterilisation process is monitored and controlled as per the equipment manufacturer’s IFU. A comprehensive equipment maintenance schedule is maintained and frequently updated by Sterilising Services.

##### Drying cabinets

Moisture from the cleaning and/or disinfection process can cause damage to the device and can adversely impact on the efficacy of disinfection and sterilisation process. Drying cabinets provide the mechanism to prevent recontamination and maintain the integrity of cleaned and/or disinfected items.

Where a drying cabinet is used, care is taken to:

not exceed temperature tolerances and ensure placement of an RMD in the cabinet allows for adequate circulation of air to support uniform drying of the device.

ensure that hot and/or compressed air used for drying a cleaned device using either a manual or an automated mechanical process is of a quality that does not increase the bioburden of the reprocessed device.

place manually cleaned RMDs below RMDs that have undergone cleaning and thermal disinfection.

a drying cabinet’s temperature is checked and recorded daily so the cabinet operates within its specified limits.

AS 5330 normative reference is referred for Drying Cabinets

##### High level disinfectors (AFERs and Antigermix)

Automated flexible endoscope reprocessors (AFERs) are used to high level disinfect semi-critical RMDs that cannot withstand moist heat sterilisation and as recommended by the manufacturer’s instructions of the RMD.

The AFERs requirements and efficacy are validated as per the ISO 15883 part 1 and part 4. High level instrument grade disinfectants that are TGA approved are used in the AFERs to disinfect the scopes.

Automated high level disinfection process is mandatory for flexible endoscopes as per the Infection Prevention and Control in Endoscopy 2021 guidelines. The quality of the final rinse water used to rinse the scopes after high label disinfection step is compliant with Table 7.3 of AS 5369 Standard.

The three micron graded filtration system is used to filter water before entering into the AFERs and is serviced every six months by the equipment manufacturer. The disinfection process records/cycle print outs that have the cycle number are recorded on the tracking system, TDOC and Digital Health Records.

Transoesophageal (TOE) probes and nasopharyngoscopes or ear, nose and throat (ENT) scopes that have no channels are high level disinfected using the UV-C device from the company Germitec called the Antigermix.

Prior to disinfection, the TOEs and the nasopharyngoscopes are pre-cleaned with detergent and water in the decontamination area. Leak testing is performed on the flexible endoscopes and ENT scopes prior to pre-cleaning. The pre-cleaned scopes are fully dried prior to disinfecting them in the Antigermix. The cycle time to disinfect TOE probes and ENT scopes in the Antigermix is 180 seconds. The Antigermix releases a sticker upon successful cycle completion which has the HLD, and the scope information printed on it to maintain traceability.

TGA approved and validated Storage bags are used to store flexible endoscopes and other RMDs for 31 days post high level disinfection, these bags are compliant with EN 16442 normative standard.

##### Validation of reprocessing equipment (installation qualification, operational qualification & performance qualification)

Prior to use, processes and the reprocessing equipment are validated by the equipment manufacturer. Most challenging (difficult to clean and sterilise due to the complex design) and multipart RMDs are validated during performance qualification. (Refer to section 7 from AS 5369)

The reprocessing equipment manufacturer should provide the following specifications prior to purchase and validation:

description of the equipment and ancillary items

a specification for the cleaning, disinfectant, or sterilising agent

a description of the instrumentation used for controlling, monitoring, and recording of cleaning, disinfecting, and sterilising processes.

the identification guide if fault recognised by the equipment.

details of safety features

the requirements for installation

a description of the software used for monitoring.

Both chemical and biological indicators are used during validation (performance qualification) of low temperature and steam sterilisation processes. Reference is made to the relevant part of ISO 15883 parts 1, 2, 3, 4, 5 and 7 for the validation of automated cleaning process. Reference is made to ISO 17665 for validation of steam sterilisation processes and ISO 22441 for validation of vaporised hydrogen peroxide sterilisation processes.

Note: According to AS 5369 clause 6.1.1, if a facility chooses to deviate from the manufacturer’s reprocessing instructions, then validation of the alternative process is undertaken and investigated in consultation with the manufacturer.

Validation of the SBS/ PSBS is undertaken in accordance with ISO 11607-2 and with reference to requirements outlined in ISO/TS 16775. If a new SBS is introduced at a time when performance qualification of the steriliser/s is not due, this process is performed independent of same if the design and the characteristics of the SBS is different.

The success of sterilisation depends on the entire reprocessing cycle and therefore all stages of the cycle are validated, including cleaning; inspection, assembly, and packaging; sterilisation; drying, storage and distribution.

Requalification is performed at least annually and whenever a change is made to either:

an RMD, its packaging material or to the cleaning, disinfection, packaging, or sterilization processes

major changes or repairs are made to reprocessing equipment that have the potential to affect the efficacy of the processes.

The responsibility for determining the necessity and extent of requalification of part or all of Installation Qualification, Operational Qualification or Performance Qualification is assigned to the facility and the decision will be made based on the risk assessment. Documented validation plans/service level agreements for the major reprocessing equipment are maintained by the departments. AS 5369 Standard Tables 10.1, 10.2 and 10.3 recommend the tests and the frequencies for each equipment type.

In situations where routine monitoring or periodic testing indicate results outside normal range or non-conformances from process parameters established during validation of the reprocessing equipment, the cause should be identified and corrected. In these situations, requalification of the equipment is considered.

Validation reports are reviewed, approved, and documented by the sterilising managers. The copies of the validation report are retained within the department for accreditation purpose.

##### Periodic preventative maintenance of reprocessing equipment including calibration of monitoring instrumentation.

Preventive maintenance of all equipment is planned and undertaken in accordance with documented procedures provided with the equipment. This is carried out by the manufacturer, their agent or a designated trained and authorized person. The procedure for every planned maintenance activity and the frequency at which the activity is to be undertaken is specified in AS 5369 Standard Tables 10.1, 10.2 and 10.3 and normative standards.

If there is an unavoidable delay to the performance of scheduled preventative maintenance, for example due to the unavailability of a qualified service person or equipment manufacturer, then the maximum timeframe that the service is delayed is determined by conducting a risk assessment and assessing the level of risk to the patient safety. The equipment manufacturer is also consulted before determining the risk.

Back to Contents

#### Section 8 - Routine monitoring and control of reprocessing equipment and processes

##### Water Quality

Water quality used in the central sterilising unit is tested to ensure it is compliant with AS 5369 Tables 7.2, 7.3 and 7.4. Water supplied is tested at the time of installation qualification to check compliance and then is tested at regular intervals as recommended in the AS 5369 standard.

Water used for manual cleaning, before final rinse, is tested every month for:

water hardness (CaCO3 equivalents) < 150mg/L

chloride < 120mg/L

by the chemical and equipment manufacturers. This is undertaken in both endoscopy and central sterilising unit.

The final rinse water quality for WD used for processing RMDs is tested monthly against AS 5369 Table 7.2 by the chemical and equipment manufacturers. If the parameters are not met then the service is isolated from use and is escalated by sterilising services. The non-compliant results and the action taken to mitigate the risk are tabled at the PCI meeting.

Steam quality tests (dryness, non-condensable gases and superheat) are not routinely performed and are tested only when issues like wet loads arise.

Feedwater for the steam sterilisers with a dedicated steam generator is tested annually by the manufacturer.

Where the results of testing deviate from the acceptance criteria in AS 5369 Table 7.2, 7.3 and 7.4 or EN 285 Table 4, corrective action shall be taken and the feedwater quality and/or steam purity retested

Various quality monitoring tests are conducted as recommended in AS 5369Tables 7.2, 7.3, 7.4, 8.1 and 8.2 to ensure RMD’s are adequately cleaned, disinfected, packaged, or sterilised and the efficacy of the process is measured and verified. The equipment manufacturer performs majority of the routine monitoring tests recommended by the ISO 15883 parts 1 and 5.

Sterilising Services performs the below checks internally to monitor the efficacy and the operation of the equipment:

Manual Cleaning process:

Protein test to test cleaning efficacy in flexible endoscopes and other RMDs (weekly)

Ultrasonic Cleaner:

Daily degas

Daily adapter check (close adapters with the tubing if not being used during the cycle)

Cavitation efficacy check (daily)

Disinfection cycle

Filter replacement (when required)

Scope buddy:

Validation and end of the day disinfection cycle (daily)

Washer disinfectors:

Check spray arms and clean filters (daily)

Check for alarms or errors (daily)

Bulk dosing system alerts

Sterilisers:

Warm up cycle (if sterilisers are idle for 4 hours or more)

Bowie & Dick test

Leak test (weekly)

Bowie & Dick function test (weekly)

Process challenge device (every cycle)

Biological indicator test (weekly)

Low temperature hydrogen peroxide sterilisers:

Leak test (monthly)

Check sterilant cup

Biological indicator test (weekly)

Heat sealers:

Seal check and ink test (daily)

Dryers:

Check the operating temperature (daily)

The outcome of all routine monitoring/tests performed are documented, including the date the test was performed, the reprocessing equipment number or the name of the RMD, the person responsible and the actual result.

##### Routine monitoring and cleaning efficacy tests for Washer disinfectors

The ISO 15883 series outline suitable methods for demonstrating the cleaning efficacy of WDs. Some of these methods are applied to manual cleaning processes (e.g. protein residue testing as described in ISO 15883‑5). An objective means of assessment is capable of detecting residues relevant to the clinical use of the RMD/other device. Testing schedule and forms for routine monitoring tests are available on the QMS.

##### Routine monitoring of flexible endoscopes / microbiological surveillance of flexible endoscopes with channels

All flexible endoscopes with channels undergo microbiological surveillance testing. Gastrointestinal endoscopes are tested in accordance with the GENCA 2025 guidelines. All bronchoscopes, duodenoscopes and EUS (endoscopic ultrasound) radial and linear are tested monthly and all colonoscopes and gastroscopes are tested quarterly (eg. cystoscopes).

Endoscopes with channels that have been reprocessed through a sterilisation cycle and stored in a wrapped state are tested at least quarterly and/or as recommended by the manufacturer’s instructions of the endoscope (eg. cystoscopes and ureteroscopes).

Microbiological surveillance is managed and performed by the sterilising services staff. Detailed guidance/work instruction on the microbiological surveillance of endoscopes and testing protocols, frequency of testing is accessible on sterilising services QMS and in Infection Prevention and Control in Endoscopy 2025 guidelines.

Back to Contents

#### Section 9 - Traceability of reprocessed semi-critical and critical RMDs

CHS Sterilising Services uses a tracking system (T-DOC) to trace semi critical and critical RMDs, including trial and loan equipment. With single instruments, only the history of the sterilisation process can be tracked on T-DOC, the cleaning history of used single instruments cannot be tracked.  Both cleaning and sterilisation handling steps/history are traceable for instrument assembled in a tray on T-DOC. The tracking system identifies the individual cycle of a specialised steriliser in which the RMD was sterilised, with documentation of parametric release or the cycle of a high-level disinfection system in which a semi critical RMD had undergone a disinfection process.

The tracking system allows identification of the reprocessing of the RMD and subsequent use on patient’s undergoing surgical or skin/mucous membrane penetration procedures.

Satellite sites in CHS that reprocesses ultrasound probes using Chronos follow a different process for tracking probes used on patients. Chronos generates a sticker to confirm that the particular probe of a certain serial number has successfully been through the high level disinfection process and is safe to be used. The information on the sticker is manually replicated on patient file or on DHR (refer to section 14).

In the event of a patient recall action, the T-DOC is used to trace an RMD/other device that is subject to numerous reprocessing cycles except for single instruments.

##### Recall of RMDs

A released RMD is recalled in a timely manner by both sterilising and perioperative services where there is evidence of failure during the cleaning, disinfection, packaging, or sterilisation processes, or during transport and storage of a reprocessed device.

In addition to the above, the RMD is recalled if there is a recall alert from TGA.

##### Handling of product complaints

if there is an issue with the sterilisation process or if a non-conforming RMD has been identified at the point of use then the product complaints or RMD non-conformances and staff accidents related to sterilising services are reported by the end user through the clinical riskman incidents and WHS safety incident report.

customer feedback/complaints, quality improvements, corrective actions and key performance indicators (KPIs) of the clinical riskman incidents/non-conformances relevant to sterilising services are discussed at the stakeholder meetings (Perioperative meetings with CH and NCH) and the PCI Standard Committee.

##### Business Continuity Plan

Appropriate systems and. a business continuity plan (BCP) is in place to ensure that the requirements of AS 5369 standards are met at all times, regardless of emergency or other suboptimal operating conditions (eg. water restrictions).

Back up procedures and the BCP are accessible on the QMS register in the sterilising services Q drive for sterilising staff only and the hardcopy is available in the reprocessing units.

This includes step by step instructions describing how to activate these procedures. The ventilation system for inspection, assembly and packaging areas and sterile storerooms is continuously operational.

Back to Contents

#### Section 10 - Audit Program

Regular audits are undertaken by Sterilising Services in areas where the RMDs used are reprocessed by sterilising services to ensure objectivity is maintained.

Frequency of audits are based on outcomes and risks involved (preferably bi-annually). Audit reports are saved in a central location and any non-conformance reported in the Perioperative Services Meetings (if applicable) and PCI Standard Committee. Non-conformance is also provided to the Quality and Safety Unit as evidence for action 3.17 of NSQHS standards.

Audits conducted in Sterilising Services include:

daily monitoring of reprocessing steps

manual cleaning processes

ultrasonic cleaners

washer disinfectors

mechanical dryers

heat sealers

steam and low temperature sterilisers

insulation testers

automated flexible endoscope reprocessors (AFERs)

Antigermix HLD

biological incubators

audits relating to the preparation of RMD trays prior to packaging.

effectiveness of the outer tray sterilisation wrap presented at the point of use, including absence of damage and aseptic presentation.

effectiveness of the tracking system is carried out by SS in collaboration with perioperative services and endoscopy department.

completion and accuracy of steam and low temperature sterilisation monitoring.

Audit reports are recorded, and the Sterilising Services summary report is tabled quarterly in the relevant Committees.

In the event of a non-conformance identified at the time of the audit, sterilising managers conduct a risk assessment, analyse the issue, develop corrective actions at the time of the audit and evaluate the controls implemented in the quality and training meeting.

The opportunity for improvement or non-conformance is recorded on the Sterilising Services Corrective Actions Register (CAR), which is tabled at the monthly Quality and Training Meeting.

The annual audit of the QMS is undertaken by an out of area auditor.

Internal audits are performed electronically using a software called as the iAuditor (IHSS initiative) where the pre-developed audit checklists/templates are available on the software. The iAuditor license is obtained from the IHSS team and the templates can be accessed on a dedicated mobile device or an iPad. Areas outside sterilising services must obtain a license and an iPad to have access to the iAuditor. The audit results can be electronically distributed to the registered email contacts through the iAuditor software.

##### Audit program for point of care reprocessing of ultrasound probes:

Managers of areas where Point of care (POC) reprocessing occurs are responsible for conducting periodic audits and ensuring audit outcomes are submitted to sterilising services, who will collate and submit to PCI Standard Committee for review and further action. Audits are to be conducted monthly, refer to Attachment 1 for audit tool.

Ultrasound probe audits

Regular audits of work practices are conducted by the area managers to identify, assess and manage the risk of transmission of infectious agents.

Internal audits are to be completed on a quarterly basis, by the area manager or by a CHS staff member who works outside the area but is familiar with the HLD process. The auditor picks a sample of random probes that were used on the patients and checks that the correct reprocessing process has been completed correctly and the process is traceable. There needs to be evidence of who completed the reprocessing (via initials or signature) and a register which contains the staff members initial/signature in the event of a follow up is required.

If there are no non-conformances identified consistently in the same audit, the audit frequency can be reduced from monthly to six monthly or from six monthly to annually if there is no mandatory requirement on the frequency of the audit.

Spot audits will occur unannounced with 24 hours’ notice by the mock accreditation assessors before an assessment is due to take place to validate the process.

Evidence required by the areas for the audit include:

point of care disinfection workflow is complying with this procedure and manufacturer’s instructions for use

user Manuals and procedures are documented

recalls and non-conformances are documented

preventative Maintenance records of the high-level disinfector are documented and monitored

validation reports including installation qualification, operational qualification, and performance qualification documented securely

annual performance requalification reports

annual electrical safety check

staff education and competency reports

traceability of reprocessed ultrasound probes

Back to Contents

#### Section 11 – Staff Education and Training

##### Sterilising Services

On commencement of employment in sterilising services, an assessment of the knowledge and competency level of new staff is undertaken.

This enables proper planning of initial and ongoing training for staff. A formal induction, orientation and training program is available to all new sterilising staff.

At a minimum, the program will include the following:

- modes of transmission of infection

- infection control principles, including standard precautions and transmission based precautions

- hand hygiene, including the need for and importance of removing jewellery, nail polish and artificial nails

- workplace health and safety issues, including the use of PPE (e.g. protective apron or gown, eye protection, closed footwear, and gloves)

- reprocessing activities

- documentation and record keeping

- an overview of quality management systems in relation to patient safety programs

On completion of training there will be an objective assessment of staff competency where the staff member demonstrates the reprocessing activities correctly to the Training and Development Manager or the Team leader.

All staff are required to participate in ongoing training and development opportunities. This includes internal and external education. All reprocessing staff require periodic reassessment of ongoing competency (annually). Where necessary, retraining will be undertaken.

Staff training and competency measures include:

new employee training program(s) and observation.

new products and processes

updates and continuing education

review of the existing training documentation and work instructions to ensure they align with the current standards

##### Clinical staff using ultrasound probes

Staff who perform reprocessing of ultrasound probes must complete the HLD training through HRIMS and be assessed competent on an annual basis. Records of the completed competency assessments are kept by Managers of the clinical staff using the HLD. Refer to Attachment 2 for Chronos Competency Assessment.

All staff who use ultrasound machines shall be trained in and have completed the following online learning modules annually, via HRIMS:

Hand hygiene

The basics on infection, prevention and control

Aseptic technique (if inserting invasive medical devices)

CHS – Hypernova Chronos Training. All users reprocessing using Chronos need to complete the eLearning package every 12 months. This eLearning is found on HRIMS.

Chronos user manual, traceability information, consumable form and network setup guidance documents are accessible on the Sterilising Services IHSS SharePoint.

Back to Contents

#### Section 12 - Ultrasound probes

The use of ultrasound is increasing in Australia within diverse care environments. Each ultrasound procedure involves contact between:

Ultrasound probe and console

Patient

Healthcare worker

Without appropriate infection control measures, ultrasound usage can put the patient and healthcare worker at risk of transmission of microorganisms and infection.

All point of care ultrasound probes need to be cleaned, disinfected, and appropriately stored after use in between multiple patient use. This includes, but is not limited to, diagnostic ultrasound machine probes, personal point of care ultrasounds (eg. butterfly iQ3 devices), bladder scanners and all types of ultrasounds used on wards (eg. dopplers).

Ultrasound probes need to be reprocessed to meet the minimum standards of practice based on the manufacturer’s IFU, Australasian Society for Ultrasound in Medicine (ASUM) guidelines, AS 5369 as well as per principles and standards outlined in Section 2.

All reusable ultrasound probes, cleaning and disinfection products and high level disinfectors or disinfection process must be approved by the Therapeutic Goods Administration (TGA) for use in healthcare. Products used for low-level disinfection must be classified as ‘instrument’ grade disinfectants. Hospital-grade or household / commercial grade disinfectants are not approved for use on RMDs, including Ultrasound probes.

The level of disinfection required is dependent on the type of procedure conducted and on patient factors.  The correct use of gel, probe covers, and correct storage of the ultrasound probes will aid in maintaining adequate disinfection, refer to Section 2.5 for more information.

The Spaulding classification system, see Table 1 in Section 1, provides a system to determine, based on intended use, the level of reprocessing for reusable medical devices, including ultrasound probes.

Because of this, the level of reprocessing required for an ultrasound probe is directly dependent on how that probe is used for each individual patient procedure. Clinical areas should consider which disinfection is required for procedures, operational considerations, such as availability of ultrasound probes, separating ultrasound probes for different indications and/or utilising a general HLD procedure. For reprocessing of ultrasound probes refer to:

Section 2.1 Non-critical

Section 2.2 Semi-critical

Section 2.4 Critical

Attachment 3 provides examples of ultrasound procedures and the required disinfection or sterilisation.

It is a requirement that the person performing the ultrasound considers the type of scan being performed, and the individual patient’s health status to ensure the correct infection prevention and control measures are used.

##### Infection prevention and control for ultrasound

Minimising the risk of transmission of pathogenic microorganisms before, during and after the Ultrasound procedure requires multiple processes in addition to the specific low- or high-level disinfection of the probe. This includes:

use of an ultrasound probe cover (sheath) to conduct the scan

type of gel (sterile versus nonsterile) used during the scan

cleaning of the ultrasound probes prior to disinfection

low- or high-level disinfection of the probe according to the type of procedure

appropriate storage of the probe after reprocessing is completed and prior to the next use

cleaning of associated equipment

traceability documentation of the procedure and successful disinfection process

- Refer to Section 2.5 for information on covers/sheaths and gel.

- Alert: If no sticker is visible on the probe, the probe must be pre-cleaned with a wipe and disinfection must be performed again before the scan can be conducted.

Back to Contents

#### Section 12.1 – Non-critical

These ultrasound probes come into contact with intact skin but not mucous membranes. Non-invasive ultrasound transducers are considered non-critical medical devices according to Spaulding’s classification and MUST undergo, as a minimum, low level disinfection (LLD) between each use. Intermediate level disinfection may be appropriate in specific circumstances.

##### Prior to using the ultrasound device:

examine ultrasound device for integrity

if not stored in the required manner perform LLD prior to use

report any defects in the probe, as any cracks may affect the cleaning and disinfection process

Alert: Cleaning is a mandatory step even if the probe was covered with a sterile cover for the procedure.

##### Procedure for low level disinfection after non-critical use:

remove and dispose of probe cover (if used)

wipe off all residual gel using paper towel or other single use product

clean/disinfect Ultrasound probe (patient end) with Clinell instrument Grade Universal wipe, or an equivalent TGA approved Medical Device Class IIb Disinfectant, moving from clean to dirty and ensuring all surfaces are dampened

separately, clean the Ultrasound handle, cord, probe holder, console, and gel bottle nozzle with another Clinell wipe.

clean Ultrasound machine screen with the manufacturer approved wipe

apply storage cover to probe or place probe in appropriate storage container.

##### Procedure for Intermediate-level disinfection

- perform hand hygiene

- wear appropriate personal protective equipment

- remove and dispose of probe cover (if used)

- wipe off all residual gel using paper towel or other single use product

- clean ultrasound probe (patient end), moving from clean to dirty and ensuring all surfaces are dampened

- dispense Tristel DUO NCU onto a dry wipe (Due wipe recommended)

- spread the foam over the surface and ensure all areas are wet

- allow the surface to remain wet for 30 seconds and leave to air dry

- separately, clean the Ultrasound handle, cord, probe holder, console, and gel bottle nozzle

- clean Ultrasound machine screen with the manufacturer approved wipe

- apply storage cover to probe or place probe in appropriate storage container

Alert: Bladder scanners and ‘doppler’ probes are currently unable to be reprocessed in the HLD machine. Therefore, a probe cover MUST be used every time a bladder scanner or Doppler Ultrasound is required to be used on a patient with:

Non-intact skin

Infected intact skin

Skin contaminated with blood or bodily fluids

A multi-resistant organism including MRSA, VRE, CPE and/or multi-resistant Gram negative

Alternatively, staff can utilise the Tristel DUO NCU

Back to Contents

#### Section 12.2 – Semi-critical

ALERT: Only staff who have been trained, and deemed competent, in performing the relevant High-Level Disinfection (HLD) process are permitted to do so. The process for deciding which HLD method is utilised is determined at a department level in conjunction with the sterilising services and infection prevention and control teams.

These Ultrasound transducers/probes  come into contact with non-intact skin, mucous membranes and blood or body fluids. These devices must undergo cleaning, followed by HLD between uses.

- Reprocessing a semi-critical device requires cleaning followed by disinfection using a HLD listed on the Australian Register of Therapeutic Goods (ARTG) and a method that is consistent with the manufacturer’s IFU.

HLD will kill all microorganisms with the exception of high numbers of bacterial endospores and is the minimum standard requirement for reprocessing ultrasound probes that are used for a semi-critical procedure. Examples include:

prior to and after use when inserting an indwelling medical device e.g. peripheral and central venous vascular catheters; pleural, ascitic or pericardial drains and/or aspiration procedures

prior to and after use when used on non-intact skin and / or mucous membranes

after having direct contact with blood and or body fluids

after use on intact infected skin (e.g. cellulitis) where NO probe cover was used

after use on intact non-infected skin in a patient requiring contact precautions for a multi-resistant organism (e.g. MRSA; VRE; MR GN) where NO probe cover was used

Alert: An automated HLD process is preferred over manual processes as there is a structured validation process to confirm that the disinfection process was completed to an adequate level. At Canberra Hospital and Weston Creek Medical Imaging, the use of  Chronos is the only method of point of care HLD. An RFID tag is programmed to work on all Chronos machines across CHS in the event one of the units fail.

##### Hypernova Chronos

The Chronos machine is an automated disinfection unit that emits high-intensity radiation in the wavelength of ‘ultraviolet type C’. CHS require the user to follow the procedure below. Staff must have completed the required training and competency assessment prior to use.

###### Process to disinfect ultrasound probe at point of care

Follow manufacturer’s IFU and refer to Germitec’ s ultrasound probe compatibility list before reprocessing.

- Don PPE (gown, gloves and goggles)

Immediately after the procedure, remove the sterile disposable cover from the transducer & dispose of in clinical waste. Change gloves

Clean the ultrasound probe with the Reynard neutral detergent wipes or tuffy wipes intended for use on medical devices to remove gel and any extraneous material from the transducer. Clinell wipes are not recommended by the HLD supplier because the chemical from the Clinell wipe tends to stick to the probe.

Alert: Cleaning is a mandatory step even if the probe was covered with a sterile cover for the procedure.

Organic residue may prevent the disinfectant from contacting all surfaces of the medical device being processed. Start the cleaning beneath the probe identification tag, cleaning follows clean to dirty by moving the wipe from part of the cable to the tip of the probe. Visually inspect the probe and repeat the process until it is visibly clean. When cleaning the probe, start wiping the cord from the electrical connector end and move towards the ultrasound transducer tip. If the wipe becomes soiled, discard and use a new wipe to clean the rest of the device

Remove gloves and dispose into waste bin, perform hand hygiene

Prepare to use the Chronos Unit. This system is a chemical free, automated, high-level disinfection process with a disinfection cycle of approximately 90 seconds. The Chronos Unit is located in the clean area

If a probe is in the unit, remove and store in appropriate storage cabinet or place in a post HLD storage and transport covers from Device Technologies (Product code: 4GMT-55PC)

Turn on Chronos Unit and associated printer. The device will beep and the word “hello” will first appear. It is ready to use when the clock time is displayed. Start up and initialisation of system will take approx. 30 seconds

When the system is ready to reprocess. Open the chamber door

Hold the probe identification tag and the cable, with the barcode facing the back of the unit. Scan the barcode tag with the barcode scanner

Load the probe into the probe cradle

The probe identification tag is detected, and the model number appears on the screen with light turning purple

Ensure the probe is above the quartz plate by 7cm

Gently close the door until you hear the locking click and ensure the probe didn’t get pinched during closure

Press the silver button, the door will lock, and the cycle will commence

The display with count from 0-100% and will declare “end” once completed (90 seconds). The cycle information will automatically print out. An electronic copy is automatically stored in the Germitrac tracking

Perform hand hygiene before taking out the disinfected probe.

Note: it is not necessary to turn the system off

The probe does not need to be wiped dry or rinsed under filtered water. There are no residues left on the probe due to UVC high level disinfection

Store probe in a protective cover and place the printed label on the cover. Probes can be returned to the console holder, but the protective cover and printed label need to remain in place until the probe is used or can be stored in a designated clean area.

Note: if an error occurs immediately notify the current distributor – Device Technologies on Technical Service: 1800 429 981 and press number 2 or techservice@device.com.au or contact Brett Houlton on bhoulton@device.com or on 0461391996. Or Charles Foster on cfoster@device.com.au or 0447211284.

The Chronos User manual is available on the Sterilising Services SharePoint site

###### Traceability for Chronos HLD

A traceability process is required for all ultrasound probes that are used for critical or semi-critical procedures. Traceability must identify the patient, the procedure and the serial number of the Ultrasound probe that was used for the procedure, in order to maintain a register that enables recall of a specific probe in case of an identified clinical risk.

The Chronos machine contains three Photodiodes within the internal chamber which records and monitors the UV-C dosage delivered to each and every cycle performed in the system. This information is then monitored and recorded on the label print outs and Germitrac software.

Records of HLD must be kept enabling recall procedures in case of decontamination failure. The Chronos system has automated electronic documentation of each cycle using the Germitrac tracking software, this can be accessed via a web browser.

The Chronos HLD has the ability to generate up to 5 stickers with all the traceability data demonstrating successful completion of the disinfection.

Traceability record

- The printer, if turned on, prints the defined number of labels after the indication panel displays “CYCLE OK”

The label/sticker is placed in the logbook against the patient details and then scanned to DHR.

For transvaginal probes only, the label or sticker is placed on the ultrasound worksheet and then is scanned this into the medical imaging section on the DHR.

The Chronos HLD generates a sticker and provides the following documentation:

Location

HLD serial or identification number

Date and time of the disinfection

Serial number of the probe if programmed

Disinfection status i.e. success or failure

There is empty space to record the name of the user and provide comments

The above information is documented on the patient files against the patient on whom the probe was used. This is either done manually by copying the information on the patient notes or by scanning the information on the sticker to DHR.

- If the Chronos is not interfaced with the ACT governments network then the information such as the name of the patient and date of birth has to be manually entered by the staff on the Chronos software.

###### Documentation of the reprocessing details of the ultrasound in the patient record

The probe details (serial number) along with the reprocessing history (disinfection process completion evidence and date performed) must be documented on the patient record.

- The information that confirms that the reprocessing has been successfully completed which is on the HLD sticker is manually entered into the patient record on DHR OR

The machine stickers that confirm that the reprocessing has been successfully completed is attached to the clinical note that consists of the patients name, MRN and date of birth and then is scanned into DHR. Some clinical areas send the traceability records to clinical records for record keeping.

There should be evidence of traceability between the RMD/probe used, patient, procedure and documentation recorded either on DHR. In the event DHR is not available, the information is recorded on a patient record as a minimum requirement.

##### Transportation for HLD

If reprocessing is not occurring at the point of care i.e. in the ward area, then the ultrasound probes are to be safely transported via a transportation trolley to ensure its safety from damage and contamination. The transport trays need to be cleaned between uses. Where washing may not be possible, the tray needs to be wiped using a disinfectant wipe.

If the tray liner system is in place where the liner is placed over the tray and then a cover ‘green for clean’ and ‘red for dirty’ is used.

When removing the tray cover and tray liner from the tray after transporting the soiled probe back to the reprocessing area, ensure cross-contamination is avoided. Remove soiled PPE (and if required don clean PPE) to clean and disinfect the tray using a neutral detergent wipe and place on the clean side of the reprocessing room, ready for the next use.

Transportation of the probes occurs between:

Point of care area to equipment storeroom after HLD – Equipment technicians or dedicated staff will put a clean disposable storage covers over the probe and transport to the storage area or cupboard in a transportation trolley.

Storage area to point of use - Equipment technicians or dedicated staff will put the probe covered in a storage cover indicating “ready for use” and safely reconnect to the appropriate ultrasound machine located by the patient bedside.

Point of use to point of care reprocessing area – The dirty ultrasound probe is safely placed back in the tray and is to be covered with a liner for clear identification of a “CONTAMINATED” probe. Safely transport and return to the Point of Care Reprocessing Area immediately for reprocessing or the HLD can be wheeled to the point of use location to disinfect the probe where possible.

##### Chronos Unit Management

Consumables required for point of care HLD using Chronos:

Thermal printer labels

Barcode tags

Barcode scanner

Cleaning wipes

Probe storage covers post high-level disinfection

Consumable order form is available on the Sterilising Services SharePoint page

###### Cleaning

External cleaning of Chronos Unit and the ring – Once a day, use a cloth or paper cloth, dampened with water or a surface cleaner/disinfectant.

Warning: Do not spread liquid inside and over the Chronos. Do not immerse the ring in any liquid.

Internal cleaning – Once a day, A super-soft cloth (optical quality) may be used to clean the quartz plate and a neutral detergent wipe to clean the outside surfaces. The remainder of the internal walls must be cleaned by Germitec Authorized Service Representative during maintenance operations.

Warning: Do not perform cleaning in the 2 minutes following a disinfection cycle. The tubes are still hot and present a burn hazard. Do not clean the tubes. Do not clean behind and in immediate proximity to the tubes.

Reprocessing Area – Clean the workbenches, reprocessing area where the HLD is kept with a surface disinfectant wipe. Sinks and the floors have to be cleaned by the external contractors (ISS) as per Infection Control recommendations.  Cleaning needs to be done daily or as per requirements and the records to confirm that the cleaning is managed by the Satellite Reprocessing Areas. Record has to be maintained by the area supervisor to indicate that it has been completed.

###### Chronos routine monitoring control

Chronos’s automated verification system replaces the manual process of checking chemical indicators and tests strips. The internal monitoring system has two photo diodes that continuously monitor the UVC dose for 90 seconds.

The Germitrac tracking is to be downloaded every month as a backup by the area supervisor if the HLD is connected to the CHS datapoint.

To download a backup:

- Go to any computer on the hospital network

Open web browser- type in IP address of Chronos (as allocated by IT department)

All cycle information is displayed in Germitrac electronic logbook

Filter for the past month, by select start date and end date

Press green button, an excel sheet of all disinfecting cycles is downloaded

Save excel file to the respective ward area’s shared drive or folder

Refer to Sterilising Services SharePoint page for detailed information from the manufacturer on Chronos Traceability Process and Chronos network set-up.

###### Replacing printer paper:

- 1. Disconnect the printer and open the top cover

2. Place the cylindrical holder into the label roll

3. Insert the end of the roll in the slot

4. Close the cover

5. Re-connect and press the power button

###### Scheduled maintenance and performance requalification

Preventative maintenance (PM) is managed by the individual areas where the Chronos machine is located. Each area organises the service agreement or a PM with the supplier of the unit i.e. Device Technologies after every 4000 cycles or after every 2 years (the unit alerts the user when nearing to 4000 cycles), whichever comes first and is part of the service contract. Only a technician approved by Germitec is authorized to perform maintenance operations on the Chronos. Preventative maintenance is undertaken in accordance with the HLD manufacturer’s instructions for use.

The preventative maintenance will include the following:

A general system check

A cleaning of external and internal parts if necessary

Changing of 6 x UV-C Light tubes

The calibration of the photodiodes

Internal chamber cleaning

Replacement of inners seals

Ventilation filter replacement

Upgrade to software if new version is available

- Alert: Only Philips UV-C tubes supplied by Germitec may be placed in the Chronos

An annual electrical safety check is undertaken by the testing and tagging team (HTM) within CHS. This is organised by the individual areas.

The supplier, Device Technologies will undertake annual Performance Qualification (validation) and provide a written or electronic report. The report will be checked by the Department Manager to ensure the machine is performing correctly or whether corrective action is required.

All validation and maintenance reports are to be saved in a secure folder by the area manager.

###### Trouble shooting

A protective quartz plate has been fitted above the bottom UVC bulbs to protect the bulbs if a probe where to be dropped.

In the event of both the quartz and UVC tubes breaking. Take system out of use:  tape with masking tape and call a Device Technologies to organise a technician to replace the bulbs and recalibrate the system.

Technical Service: 1800 429 981 or techservice@device.com.au

Refer to Sterilising Services SharePoint page for Chronos user manual.

Back to Contents

#### Section 12.3 – Critical

These items confer a high risk for infection if they are contaminated with any microorganism and must be sterile at the time of use. Ultrasound probes which have been used for a critical procedure require reprocessing via sterilisation. Ultrasound probes which have been used in a sterile area of the body are an example of an ultrasound probe which requires sterilisation. Sterilisation must be performed in the sterilising unit by a sterilisation method that is recommended by the manufacturer’s instructions for use.

ALERT: Most ultrasound probes are heat sensitive and are unable to be sterilised using methods that involve high temperatures. Therefore, the reprocessing requirements for probes that are heat resistant and are used for a critical procedure are reprocessed by Sterilising Services.

##### Transportation of Ultrasound Probes

When transporting a reprocessed ultrasound probe to the point of care, or returning it to the reprocessing area, it is important that the probe remains covered in a sealed storage cover or in an enclosed transport container to prevent environmental and probe contamination.

Transport containers will require wiping with a neutral detergent wipe after the probe is removed and before returning the tray to the storage area.

Additionally, the probe holder on the ultrasound machine must be cleaned, prior to returning the sterilised probe to the ultrasound machine.

##### Probe Storage

Appropriate ultrasound probe storage is required once cleaning and sterilisation has been completed, to protect from environmental contamination whilst not in use.  An ultrasound probe which has undergone sterilisation in the sterilising services unit must have a TDOC label (sterilising services traceability system that generates a label with the reprocessing details) attached to the probe with reprocessing details as evidence of appropriate reprocessing.

Back to Contents

#### Section 12.4 - Ultrasound Gel Selection, Probe Covers and Probe Storage

##### Ultrasound Transmission Gel

Ultrasound transmission gel is a viscous gel that acts as a contact medium between the patient and the ultrasound probe, to enable the acoustic transmission of ultrasound waves. The use of sterile or non-sterile Ultrasound gel depends on the type of procedure being performed.

There are two types of Ultrasound gel at CHS:

###### Non-sterile gel

The use of non-sterile gel should be limited to low risk / non-critical ultrasound scans i.e. those conducted on intact skin.

- 250mL Multiuse bottles

Check expiry date prior to each use

The date the bottle was opened must be clearly marked on the bottle

Must be discarded once empty and should not be refilled

The bottle should be sealed when not in use

Should not have direct contact between the dispensing tip and any person or transducer

If there is evidence of contamination or damage to the package integrity, any unused portions must be discarded

Must be discarded 28 days after opening.

###### Sterile gel

Sterile gel shall be used for all semi-critical and critical ultrasound scans.

Single use sachet 20g

Must be clearly marked as sterile and must be unopened prior to use

Once opened unused portions must be discarded

###### Gel warming

The heating or warming of gel is not recommended due to the risk of bacterial contamination and growth. However, in circumstances where warm gel is necessary, the use of dry heat is the preferred method. Gel warmers must be cleaned and disinfected regularly in accordance with the manufacturer’s instructions for use. For multiuse bottles, the 28 day discarding advice must be strictly adhered to.

###### Gel Storage

Non-sterile gel can be stored in any appropriate storage cupboard, as well as on the side of the ultrasound machine.

Sterile gel should be stored in an appropriate sterile stock area as per the manufacturer’s recommendations for storage of the consumable.

##### Probe Covers

Probe covers are recommended for all ultrasound scans assessed as being semi-critical or critical, including on intact infected skin and intact non-infected skin in a patient colonised with an MRO. Covers are not required for non-critical tasks on intact healthy skin. Probe covers are available as sterile and non-sterile. Staff should ensure they are choosing the correct cover for the procedure they are undertaking.

Alert: Probe covers can be perforated or contain small, unrecognised defects. As a result, HLD is still required even if a sterile cover is used for activities that require high-level disinfection.

##### Probe Storage

Appropriate ultrasound probe storage is required once cleaning and disinfection has been completed, to protect from environmental contamination whilst not in use.  An ultrasound probe which has undergone HLD must have evidence of appropriate reprocessing by a HLD sticker being on the probe/storage cover.

Alert: If no sticker is visible on the device, the probe must be pre-cleaned with a wipe and HLD must be performed before the scan can be conducted.

Ultrasound probes need to be stored in an appropriate environment to protect from environmental contamination.  A specific cabinet is recommended but if not available the minimum standard recommended is a clean disposable cover applied to the probe to mitigate risks from environmental contaminants. In a point of care setting, a post HLD probe storage covers with tamper proof seal is recommended.

Note: Remove the storage and transport covers before performing the procedure.

Where probes are stored in a specific storage trolley with open trays, the probe should be placed in a disinfected plastic tray and then covered with an elasticised plastic cover. Evidence of appropriate reprocessing by a HLD sticker being on the probe/storage tray must be visible.

Alert: The storage cover on the probe cannot be used as the probe cover for the scan, nor can it be placed back on the probe after its use. It is to be removed and discarded prior to use.

##### Back to Contents

#### Section 12.5 – Acquisition of new Ultrasound machines, probes and consumables

To ensure all future purchases of ultrasound machines and probes comply with this policy, all ultrasound probes must be compatible with the current CHS HLD manufacturer’s instructions and warranty requirements prior to purchase.

When new or replacement ultrasound machines, ultrasound probes, and/or associated consumables are purchased, these must be assessed according to the purchase of RMDs, reprocessing agents, reprocessing equipment and clinical consumables processes already available across CHS.

ALERT: Personal ultrasound devices must be approved by and registered with the Healthcare Technology Management (HTM) asset management team. The staff member must outline the method of HLD for the ultrasound probe when registering the device.

##### Back to Contents

#### Evaluation of the policy and processes

Outcome

RMD’s are managed as specified in this policy

Ultrasound probes are managed and disinfected as per the policy

Measures

regular audits are undertaken to review activities and ensure quality outcomes. Audit results are reported to the PCI Standard Committee, quarterly

ISO 9001:2015 Quality Management System annual follow up audits are undertaken by the external auditors to validate the existing processes and the QMS in accordance with the requirements of the ISO 9001:2015 QMS standard

reprocessing activity/productivity and clinical riskman incident reporting KPI’s relevant to sterilising services are reported on the monthly Dashboard/Scorecard and are tabled at relevant meetings

Back to Contents

#### Related policies, procedures, guidelines and legislation

##### Frameworks

CHS Corporate Plan

CHS Strategic Plan

##### Policies

Waste Management

Incident Management – Clinical

Clinical Records Management

##### Procedures

Infection Prevention and Control

Aseptic Technique

Incident Management – Clinical

Management of Recalls, Alerts and Product Corrections

Management of Occupational Blood and Body Fluid Exposures (BBFE) Procedure

##### Guidelines

Infection Control – Cleaning (Shared) Patient Care Equipment

##### Legislation

Government Procurement Act 2001

Health Records (Privacy and Access) Act 1997

Human Rights Act 2004

Work Health and Safety Act 2011

Carers Recognition Act 2021

##### Standards

National Safety and Quality Health Service (NSQHS) Standards

Australian Standard AS 5369:2023 Reprocessing of reusable medical devices and other devices in health and non-health related facilities

Australasian Health Facility Guidelines (Sterilizing Services and Endoscope Reprocessing Unit B.0190)

Infection Prevention and Control in Endoscopy 2025

##### Other

Australian Charter of Healthcare Rights

The Australian Guidelines for the Prevention and Control of Infection in Healthcare

Guidelines for reprocessing ultrasound transducers. Australas J Ultrasound Med 2017; 20 (1): 30–40. Basseal JM, Westerway S, Juraja M, et al

Recommendations from the College of Intensive Care Medicine Ultrasound Special Interest Group (USIG)- AJUM May 2020 23 (2) Costello et al.

Australasian Society for Ultrasound in Medicine guidelines

Back to Contents

#### References

- National Infection Control Guidelines specific to Creutzfeldt-Jakob disease at: https://www.health.gov.au/sites/default/files/documents/2022/05/cjd-infection-control-guidelines.pdf

Australian Commission on Safety and Quality in Health Care. National Safety and Quality Health Service Standards.

https://www.safetyandquality.gov.au/sites/default/files/migrated/National-Safety-and-Quality-Health-Service-Standards-second-edition.pdf

https://www.safetyandquality.gov.au/our-work/infection-prevention-and-control/australian-guidelines-prevention-and-control-infection-healthcare

National Safety and Quality Health Service (NSQHS) Standards- Advisory note version 10 AS18/07: Reprocessing of reusable medical devices in health service organisations https://www.safetyandquality.gov.au/standards/nsqhs-standards/assessment-nsqhs-standards/nsqhs-standards-advisories/advisory-as1807-reprocessing-reusable-medical-devices-health-service-organisations

Normative References from AS 5369 Standard. (Multiple standards within AS 5369:2023)

Infection Prevention and Control in Endoscopy 2025 (Gastroenterological Nurses College of Australia – GENCA) https://www.genca.org/

Transoesophageal Echocardiogram https://www.saheart.com.au/services/diagnostic-tests/transoesophageal-echocardiogram-(toe).html

The Australasian Health Facility Guidelines (AusHFG). 0190 – Sterilising Services and Endoscope Reprocessing Units; https://healthfacilityguidelines.com.au/news/aushfg-content-update-sterilizing-services-unit-hpu

Australian Guidelines for the Prevention and Control of Infection in Healthcare (NHMRC). https://www.safetyandquality.gov.au/sites/default/files/202209/australian_guidelines_for_the_prevention_and_control_of_infection_in_health_care_-_current_version_-_v11.13_19_september_2022.pdf

TGO 104 https://www.legislation.gov.au/Details/F2019L00482

SS Decontamination, packaging, and sterilisation procedures on the QMS. (STR-GEN-PR-116, STR-GEN-PR-118 and STR-GEN-PR-117).

https://www.health.gov.au/sites/default/files/documents/2022/05/cjd-infection-control-guidelines.pdf

Australian College of Perioperative Nurses (ACORN). 2023. Standards for Safe and Quality Care in the Perioperative Environment (SSQCPE) for Organisations

Sterilising Services and Endoscope Reprocessing Unit - B.0190 [Accessed 22 December 2023]

https://onlinelibrary.wiley.com/doi/full/10.1002/ajum.12042

https://healthfacilityguidelines.com.au/hpu/sterilizing-services-and-endoscope-reprocessing-unit

Advisory note DI21/04; Disinfection requirements for reusable semi critical medical devices - Standard 1.6: Healthcare Associated Infection (under review)

Australasian Sonographer Association. ASA Clinical Statement, the safe use and storage of ultrasound gel (2021).

Australian Commission on Safety and Quality in Health Care. Disinfection requirements for reusable semi critical medical devices – Standard 1.6: Healthcare Associated Infection.  Advisory DI21/04.

Australian Commission on Safety and Quality in Health Care. Cleaning and disinfection of ultrasound transducers. Fact Sheet. Accessed 25 March 2025.

Australian Guidelines for the Prevention and Control of Infection in Healthcare, Canberra: National Health and Medical Research Council (2019).

Australian Register for Therapeutic Goods https://www.tga.gov.au/australian-register-therapeutic-goods

Australia/New Zealand Standard: AS5369:2023 Reprocessing of Reusable Medical Devices and other devices in health and non-health related facilities. Standards Australia

Basseal, J. M., Westerway, S. C., Juraja, M., van de Mortel, T., McAuley, T. E., Rippey, J., & Mizia, K. (2017). Guidelines for reprocessing ultrasound transducers. Australian J Ultrasound Med, 20(1), 30-40.

Costello, C., Basseal, J. M., Yang, Y., Anstey, J., & Yastrebov, K. (2020). Prevention of pathogen transmission during ultrasound use in the Intensive Care Unit: Recommendations from the College of Intensive Care Medicine Ultrasound Special Interest Group (USIG). Australasian Journal of Ultrasound in Medicine.

Diagnostic Imaging Accreditation Scheme (DIAS). (2015). Department of Health, AustralianGovernment.http://www.commcarelink.health.gov.au/internet/main/publishing.nsf/Content/practice-accreditation-standards-jan2016

Ferhi, K., Rouprêt, M., Mozer, P., Ploussard, G., Haertig, A., & de La Taille, A. (2013) Hepatitis C transmission after prostate biopsy. Case reports in urology, 2013.

Hiyama, Y., Takahashi, S., Uehara, T., Ichihara, K., Hashimoto, J., & Masumori, N. (2016). A case of infective endocarditis and pyogenic spondylitis after transrectal ultrasound guided prostate biopsy. Journal of Infection and Chemotherapy, 22(11), 767-769.

Oleszkowicz, S. C., Chittick, P., Russo, V., Keller, P., Sims, M., & Band, J. (2012). Infections associated with use of ultrasound transmission gel: proposed guidelines to minimize risk. Infection Control & Hospital Epidemiology, 33(12), 1235-1237.

Storment, J. M., Monga, M., & Blanco, J. D. (1997). Ineffectiveness of latex condoms in preventing contamination of the transvaginal ultrasound transducer head. Southern medical journal, 90(2), 206-208.

Westerway, S. C., Basseal, J. M., Brockway, A., Hyett, J. A., & Carter, D. A. (2017). Potential infection control risks associated with ultrasound equipment–a bacterial perspective. Ultrasound in medicine & biology, 43(2), 421-426.

the Australasian Society for Ultrasound in Medicine and the Australian Sonographer’s Association and the Australasian College for Infection Prevention and Control (2017)

Standards EN 14561 , EN 14562 , EN 14563 and EN 14476 adapted to Chronos

AOAC use-dilution methods : 955.14, 955.15, 955.17, 964.02. ASTM E 1053-97

https://germitectraining.com/lessons

Chronos User Manuals - Google Drive

Referred to the Government of Western Australia, WA Country Health Services template

- Back to Contents

#### Definition of terms

Cleaning: The removal of contamination from an item to the extent necessary for further reprocessing or for intended use.

Disinfection: The destruction of any microorganisms using thermal or chemical means.

- Low level instrument grade (LLD): destroys all vegetative bacteria (except tubercle bacilli), lipid viruses, some nonlipid viruses, and some fungi, but not bacterial spores.

- Intermediate level instrument grade (ILD): kills all microbial pathogens, except bacterial endospores, and is bactericidal, fungicidal (not necessarily chlamydospores or sexual spores), and virucidal.

High level instrument grade (HLD): kills all microbial pathogens, except large numbers of bacterial endospores. Medical device: Is an instrument, apparatus, appliance, material or other article intended for: diagnostic, prevention, monitoring, treatment and/or alleviation of disease.

Reprocessing: All activities required to ensure that a used RMD is safe for its intended purpose. NOTE: This is a multi-step process that includes cleaning, disinfection, sterilisation (if applicable), and appropriate storage.

Reusable Medical Device (RMD): A medical device that is designed or intended by its manufacturer as suitable for reprocessing and reuse for multiple patients. It is not a medical device that is designated or intended by its manufacturer for single use only.

Spaulding classification Scheme: Is a strategy for guiding on the reprocessing methods to undertake for contaminated medical devices to deem then safe for its intended use. The system classifies a reusable medical device based on the risk to patient safety from contamination after the procedure: Noncritical, semi-critical or critical.  The system also established three levels of microbiocidal methods/processes with the three classes of medical devices: Low-level disinfection, high-level disinfection, and sterilisation.

Sterilisation: Destroys microorganisms on an object rendering it free from viable microorganisms

Back to Contents

#### Search terms

Sterilising Services; Reusable medical devices; Reprocessing; AS 5369; Sterilisation; RMDs; purchase, Manufacturer’s Instructions for Use, GENCA, ultrasound probes, Chronos, Tristel, traceability, high level disinfection, semi critical, HLD, Chronos

- Back to Contents

#### Attachments

Attachment 1 - Audit tool for Reprocessing of ultrasound probes at point of care

Attachment 2 - High Level Disinfection of Ultrasound Probes using Chronos Competency Assessment

Attachment 3 - Examples of Ultrasound Cleaning and Reprocessing

Back to Contents

For Policy Team to complete:

| Date amended | Section amended | Divisional approval | Final approval |
| --- | --- | --- | --- |
| 15 December 2025 | Complete review | Chris Tarbuck, EBM, IHSS | CHS Policy Document Review Panel |
| 29 January 2026 | Re-formatted Attachment 1 to meet accessibility requirements. | CHS Policy and Consumer Handouts Team | CHS Policy and Consumer Handouts Team |

This document supersedes the following:

| Document number | Document name |
| --- | --- |
| CHS24/643 | Reprocessing of Reusable Medical Devices |

Disclaimer

This document has been developed by Canberra Health Services specifically for its own use.  Use of this document and any reliance on the information contained therein by any third party is at his or her own risk and Canberra Health Services assumes no responsibility whatsoever.

- Back to Contents

| Acknowledgement of Country Canberra Health Services acknowledges the Ngunnawal people as traditional custodians of the ACT and recognises any other people or families with connection to the lands of the ACT and region. We acknowledge and respect their continuing culture and contribution to the life of this region. © Australian Capital Territory, Canberra 2024 | Accessibility call (02) 5124 0000 Interpreter call 131 450 canberrahealthservices.act.gov.au/accessibility |
| --- | --- |

#### Attachment 1 – Reprocessing of Ultrasound Probes at Point of Care Internal Audit Form

| Audit location/area: | Audited by: | Audit Date: |
| --- | --- | --- |
| Specialty of the unit: | Manager of the area: | Staff audited (if applicable): |

Note: CNC’s or Sonographers of individual ward areas have to ensure that the audit is undertaken every month, and the audit results have to be submitted to the PCI Committee.

|  |  | Was the task completed | Was the task completed? | Was the talk completed? |  |
| --- | --- | --- | --- | --- | --- |
| No. | Criteria | Yes | No | Not applicable | If not compliant, the action taken/comments |
|  | Management Systems |  |  |  |  |
| 1 | Are staff undertaking disinfection process familiar with the RMD policy and documented procedure for reprocessing of ultrasound probes? |  |  |  |  |
| 2 | Are the manufacturer’s instructions or user manual available for the High-Level Disinfection system (eg. Chronos) in your area? |  |  |  |  |
| 3 | Does your area keep a record of the manufacturer’s instructions for the different model numbers of the ultrasound transducers that are used? |  |  |  |  |
| 4 | Have staff performing high level disinfection completed the online training? |  |  |  |  |
| 5 | Have staff performing high level disinfection process completed competency assessment in the past 12 months? |  |  |  |  |
| 6 | Is the HLD machine due for a preventative maintenance or service? Is the service report available with your area manager? (service is undertaken once every 2 years or 4000 cycles) |  |  |  |  |
| 7 | Has the HLD machine undergone Performance Requalification (validation) within the last 12 months & is the report available in your area? (requirement is annual) |  |  |  |  |
| 8 | Are the ultrasound probes attached to the correct RFID/barcode tags? Are the tags programmed with the correct serial number as on the probe? |  |  |  |  |
| 9 | Is there a process to identify & track the disinfection history of a probe that was used on the patient? |  |  |  |  |
| 10 | Can you successfully retrieve reprocessing information for a reprocessed ultrasound probe in the event of a recall? |  |  |  |  |
|  | Clinical Practice |  |  |  |  |
| 1 | Is a disposable sterile sheath or cover used with the ultrasound probe between each use? |  |  |  |  |
| 2 | (Semi-critical probes or probes that may come into close proximity with broken skin / wounds) |  |  |  |  |
| 3 | Do the ultrasound probes get disinfected between each use? |  |  |  |  |
| 4 | Were the ultrasound probes cleaned prior to high level disinfection? |  |  |  |  |
| 5 | Are the cleaning wipes used compatible with the High Level Disinfection process? |  |  |  |  |
| 6 | Are staff involved in cleaning of the probe wearing the correct PPE? |  |  |  |  |
| 7 | Are staff following the unidirectional workflow i.e dirty to clean and ensure there is no overlap or cross contamination between dirty, clean, and disinfected probes? |  |  |  |  |
| 8 | Are staff changing gloves in between dirty and clean steps? |  |  |  |  |
| 9 | Are you using the correct method and process to disinfect ultrasound probes? (check manufacturer’s instructions of US probes) |  |  |  |  |
| 10 | Are reprocessed probes correctly covered with storage covers? (If being used) |  |  |  |  |
| 11 | Are reprocessed probes easily distinguishable from unprocessed ones? |  |  |  |  |
| 12 | Are labels/stickers generated correctly attached to the probes as evidence? (High-Level Disinfection tracking label/sticker) |  |  |  |  |
| 13 | Are the disinfected probes being stored in appropriate storage conditions after high level disinfection? |  |  |  |  |
| 14 | Are you aware that the storage covers for storing the probes after disinfection are not for use on the patients? |  |  |  |  |

Note: Audit findings are required to be reported to the PCI Standard Committee meeting.

Additional comments, if any:

#### Attachment 2 – High Level Disinfection of Ultrasound Probes using Chronos Competency Assessment

| Staff Name: |  | Date Assessed: |  |
| --- | --- | --- | --- |
| Area/Department: |  | Assessed by/Assessor: |  |

| Pre-Clean Point of Care (if assessing clinician) | Y | N |
| --- | --- | --- |
| Remove and dispose of probe cover (if used) |  |  |
| Wipe off residual gel with low-linting cloth |  |  |
| Wipe probe with a recommended instrument grade neutral wipe approved by the HLD manufacturer |  |  |
| Perform hand hygiene and apply clean gloves |  |  |
| Use a fresh, second instrument grade neutral wipe for the probe handle, cord and console. If applicable, continue to the probe holder on the machine |  |  |
| Clean screen with a manufacturer approved wipe |  |  |
| Remove gloves and perform hand hygiene |  |  |
| Comments: |  |  |

| Transport to Reprocessing Area |  | Y | N |
| --- | --- | --- | --- |
| Ultrasound probe is placed in probe holder on machine, ultrasound machine is wheeled to reprocessing area, OR the probe is disconnected and taken to reprocessing area |  |  |  |
| Perform hand hygiene |  |  |  |
| Comments |  |  |  |
| Reprocessing/High Level Disinfection |  | Y | N |
| Open door and hang probe inside by the barcode tag (flat side facing the reader) |  |  |  |
| Loosely thread the cable out of the door so as not to pull out of cradle |  |  |  |
| Put plug on top of unit if plugged into ultrasound machine |  |  |  |
| Check probe detection light is purple |  |  |  |
| Press the silver button |  |  |  |
| Remove gloves and perform hand hygiene |  |  |  |
| When cycle is complete, blue light is on |  |  |  |
| Remove probe from Chronos by holding barcode tag |  |  |  |
| Place POST HLD cover, with tamper sticker |  |  |  |
| Place cycle sticker/label on cover |  |  |  |
| Comments: |  |  |  |
| Transport / Storage |  | Y | N |
| Ultrasound Probes: |  |  |  |
| Transport single ultrasound probe (with storage cover) to machine and hang in probe hanger OR place in designated hanging rack or storage drawer |  |  |  |
| Transport probe to the destination |  |  |  |
| Comments |  |  |  |
| Documentation |  | Y | N |
| Place the HLD sticker/label on the patient file or scan and save or hand write the details on the sticker on the patient notes in DHR |  |  |  |
| Record reprocessing information such as serial number of the probe in patient files |  |  |  |
| Comments |  |  |  |
| Overall Assessment rating: | C | NC | RD |

C = competent, NC = not competent, RD = requires development.

#### Attachment 3 – Examples of Ultrasound Cleaning and Reprocessing

The following table is designed to provide staff with an example of:

the type of ultrasound being performed

the required gel and probe cover for use

the required cleaning, reprocessing and storage of the probe

the required traceability documentation

This is not an exhaustive list. If staff are uncertain, they should speak with their manager or contact the infection prevention and control unit, prior to use.

Areas, particularly in large areas with many staff, that have ultrasounds which are used interchangeably for both non-critical and critical activities may choose to routine use high level disinfection for consistency and to avoid potential errors.

| Activity Type | Probe has contact with | Example of ultrasound being performed | Probe cover | Gel | Cleaning of non probe equipment | Reprocessing | Storage Cover | Traceability document |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Non- Critical | Intact, non-infected skin only AND Patient is NOT colonised with a multi-resistant organism (MRO) AND No open wound or body fluid exposure | Surface ultrasound (e.g. abdomen or bladder scan) | Not required | Single use or multi-use, non-sterile | Required after every use | Low level disinfection | Yes | Not Required |
| Non-critical | Infected intact skin OR Patient is colonised with MRO | Surface ultrasound on patient with an MRO (e.g. VRE) OR Surface ultrasound involving an area of cellulitis | Required, single use, non-sterile | Single use or multi-use, non-sterile | Required after every use | Low level disinfection | Yes | Not Required |
| Semi critical | Non-intact skin and/or mucous membranes OR Ultrasound probes that have had direct contact with blood/body fluids (ILD is suitable if probe not contaminated with blood) OR Ultrasound probes used for the insertion of indwelling devices | Surface ultrasound on open wounds or unhealed surgical wounds OR Ultrasound guided vascular access devices (e.g. PIVC, PICC, CVC, arterial lines) (ILD is suitable if probe not contaminated with blood) OR Ultrasound guided insertion of drains OR Transvaginal or transrectal Ultrasound | Required, single use, sterile | Single use sterile | Required after every use | High level disinfection | Sterile or vendor provided probe storage cover | Required |
| Critical | Sterile areas of the body | Intraoperative procedures under Ultrasound guidance OR Ultrasound guided biopsies | Required, single use, sterile | Single use sterile | Required after every use | Sterilisation | Sterile or vendor provided probe storage cover | Required |

- Back to Contents